## Guidance
Double-Click Icon to Open Guidance Document
Then proceed to DAP Applicable Requirements worksheet (tab below)

## DAP Applicable Requirements
DAP Applicable Requirements - ISO/IEC 17025:2017 with cross-reference to 2005 |  |  |  |  | Client/Customer Support Guidance-ULID-000827 (DCS:00-OP-C0043) v10.1
Please reference the Guidance Worksheet of this document for information on how to read this DAP Applicable Requirements table.
This table identifies the requirements that apply for the DAP program, breaking down the applicability of requirements by whether a lab is accredited, or non-accredited, and whether they perform any calibration in-house.
Requirements highlighted in yellow are new per the OSHA Directive CPL 01-00-004 - These requirements apply only to CTDP/TCP and TPTDP/PPP laboratories
<-- Press "-" to minimize or "+" to expand Section 1. Numbers 1 or 2 in the left Margin will do the same.
SECTION 1 - Administrative & Agreement Items - The following administrative and agreement items will be assessed every year.
Requirement | Requirement Text |  | Additional Notes | Lab Applicability | Remarks
Administrative | If DAP participant lab is accredited the UL Assessor will confirm accreditation on Accreditor's website |  |  | Accredited | UL Assessor will go to accreditor's site and pull up lab's accreditation. Confirm address, management and dates of accreditation coverage.
Administrative | If lab is accredited, the UL Assessor will confirm that the accreditor is a signatory under ILAC, EAC or APLAC |  |  | Accredited | While on the accreditor's site, accreditor will confirm their affiliation. If the lab is shown to be accredited by an accreditor that is a signatory to ILAC, EAC or APLAC, the management system assessment scope will follow the "Accredited" audit track.
Administrative | Confirm the correct agreement is in place between UL and the DAP participant laboratory. |  |  | Accredited & Non-Accredited | Assessor will confirm the correct agreement is in place for the DAP participant's program. CTDP: L56 TPTDP: L53, Third Party by Invitation (TPI), modified L56 or Subcontractor Services Agreement (SSA) (SSA will not be in ECM - no need to confirm) Confirm that agreement is signed by the customer.
Agreement | Per clause 1.2a of most DAP agreements - the customer must maintain a list of the standards and tests for which the test facility has been qualified. |  | DAP Participants are responsible for maintaining their own scope of DAP participation and only submitting test data according to that scope of participation. | Accredited & Non-Accredited | Assessor will confirm that the customer maintains a DAP scope and understands it is their responsibility to limit their DAP submissions to data from those tests.
Agreement | Confirm Confidentiality is addressed per clause 6.2 of UL's Global Supplier Code of Business Conduct, found at www.ul.com/ulglobalethics. |  | Suppliers are expected to safeguard UL’s confidential information and that of its clients by keeping it secure and limiting access to those who need to know to do their job. | Accredited & Non-Accredited | TPTDP ONLY: How does the lab address confidentiality? Protect customer/UL data? Assure staff keeps information confidential?
Agreement | Confirm Conflicts of Interest are addressed per clause 6.3 of UL's Global Supplier Code of Business Conduct, found at www.ul.com/ulglobalethics. |  | Suppliers must immediately disclose to UL Management actual or potential conflicts of interest, including any business relationship and/or any financial interest of a UL employee in a Supplier's business | Accredited & Non-Accredited | Assessor will confirm that the lab has a process to identify and address conflicts of interest and are aware of the requirement to communicate them to UL.
<-- Press "-" to minimize or "+" to expand Section 1. Numbers 1 or 2 in the left Margin will do the same.
SECTION 2 - ISO/IEC 17025 Applicable Requirements
2017 Clause | Standard Text (ISO/IEC 17025:2017 or OSHA Directive:2019) | 2005 Clause # | ISO/IEC 17025:2005 Text | Lab Applicability | Remarks
4 | General Requirements
4.1.1 | Impartiality
4.1.1 | Laboratory activities shall be undertaken impartially and structured and managed so as to safeguard impartiality. | 4.1.5 d | The laboratory shall have policies and procedures to avoid involvement in any activities that would diminish confidence in its competence, impartiality, judgement or operational integrity; | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
4.1.2 | The laboratory management shall be committed to impartiality. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
4.1.3 | The laboratory shall be responsible for the impartiality of its laboratory activities and shall not allow commercial, financial or other pressures to compromise impartiality. | 4.1.5 b | The laboratory shall have arrangements to ensure that its management and personnel are free from any undue internal and external commercial, financial and other pressures and influences that may adversely affect the quality of their work; | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
4.1.4 | The laboratory shall identify risks to its impartiality on an on-going basis. This shall include those risks that arise from its activities, or from its relationships, or from the relationships of its personnel. However, such relationships do not necessarily present a laboratory with a risk to impartiality. NOTE A relationship that threatens the impartiality of the laboratory can be based on ownership, governance, management, personnel, shared resources, finances, contracts, marketing (including branding), and payment of a sales commission or other inducement for the referral of new customers, etc. | 4.1.4 | If the laboratory is part of an organization performing activities other than testing and/or calibration, the responsibilities of key personnel in the organization that have an involvement or influence on the testing and/or calibration activities of the laboratory shall be defined in order to identify potential conflicts of interest. NOTE 1 Where a laboratory is part of a larger organization, the organizational arrangements should be such that departments having conflicting interests, such as production, commercial marketing or financing do not adversely influence the laboratory's compliance with the requirements of this handbook. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
4.1.5 | If a risk to impartiality is identified, the laboratory shall be able to demonstrate how it eliminates or minimizes such risk. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
4.2 | Confidentiality
4.2.1 | The laboratory shall be responsible, through legally enforceable commitments, for the management of all information obtained or created during the performance of laboratory activities. The laboratory shall inform the customer in advance, of the information it intends to place in the public domain. Except for information that the customer makes publicly available, or when agreed between the laboratory and the customer (e.g. for the purpose of responding to complaints), all other information is considered proprietary information and shall be regarded as confidential. | 4.1.5 c | The laboratory shall have policies and procedures to ensure the protection of its customers' confidential information and proprietary rights, including procedures for protecting the electronic storage and transmission of results | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
4.2.2 | When the laboratory is required by law or authorized by contractual arrangements to release confidential information, the customer or individual concerned shall, unless prohibited by law, be notified of the information provided. |  |  | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
4.2.3 | Information about the customer obtained from sources other than the customer (e.g. complainant, regulators) shall be confidential between the customer and the laboratory. The provider (source) of this information shall be confidential to the laboratory and shall not be shared with the customer, unless agreed by the source. |  |  | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
4.2.4 | Personnel, including any committee members, contractors, personnel of external bodies, or individuals acting on the laboratory’s behalf, shall keep confidential all information obtained or created during the performance of laboratory activities, except as required by law. |  |  | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
5 | Structural requirements
5.1 | The laboratory shall be a legal entity, or a defined part of a legal entity, that is legally responsible for its laboratory activities. NOTE For the purposes of this document, a governmental laboratory is deemed to be a legal entity on the basis of its governmental status. | 4.1.1 | The laboratory or the organization of which it is part shall be an entity that can be held legally responsible. | No | The organizational issues are addressed by the L56 (CTDP agreement) or the appropriate TPTDP agreement. It covers the complete scope and associated limitations of the testing and test data submitted to UL
5.2 | The laboratory shall identify management that has overall responsibility for the laboratory | 4.1.5 h | The laboratory shall have technical management which has overall responsibility for the technical operations and the provision of the resources needed to ensure the required quality of laboratory operations; | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.3 | The laboratory shall define and document the range of laboratory activities for which it conforms with this document. The laboratory shall only claim conformity with this document for this range of laboratory activities, which excludes externally provided laboratory activities on an ongoing basis. |  |  | No | The organizational issues are addressed by the L56 (CTDP agreement) or the appropriate TPTDP agreement. It covers the complete scope and associated limitations of the testing and test data submitted to UL
5.4 | Laboratory activities shall be carried out in such a way as to meet the requirements of this document, the laboratory’s customers, regulatory authorities and organizations providing recognition. | 4.1.2 | It is the responsibility of the laboratory to carry out its testing and calibration activities in such a way as to meet the requirements of this handbook and to satisfy the needs of the customer, the regulatory authorities or organizations providing recognition. | No | The organizational issues are addressed by the L56 (CTDP agreement) or the appropriate TPTDP agreement. It covers the complete scope and associated limitations of the testing and test data submitted to UL
5.4 | This shall include laboratory activities performed in all its permanent facilities, at sites away from its permanent facilities, in associated temporary or mobile facilities or at a customer's facility. | 4.1.3 | The management system shall cover work carried out in the laboratory’s permanent facilities, at sites away from its permanent facilities, or in associated temporary or mobile facilities. | No | The organizational issues are addressed by the L-56 that CTDP and TCP participants are required to sign. It covers the complete scope and associated limitations of the testing and test data submitted to UL
5.5 | The laboratory shall: |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.5 a | a) define the organization and management structure of the laboratory, its place in any parent organization, and the relationships between management, technical operations and support services. | 4.1.5 e | The laboratory shall define the organization and management structure of the laboratory, its place in any parent organization, and the relationships between quality management, technical operations and support services; |  | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.5 b | b) specify the responsibility, authority and interrelationship of all personnel who manage, perform or verify work affecting the results of laboratory activities | 4.1.5 f | The laboratory shall specify the responsibility, authority and interrelationships of all personnel who manage, perform or verify work affecting the quality of the tests and/or calibrations; |  | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.5 c | c) document its procedures to the extent necessary to ensure the consistent application of its laboratory activities and the validity of the results | 4.2.1 | The laboratory shall document its policies, systems, programs, procedures and instructions to the extent necessary to assure the quality of the test and/or calibration results. |  | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.6 a-e | The laboratory shall have personnel who, irrespective of other responsibilities, have the authority and resources needed to carry out their duties, including: a) implementation, maintenance and improvement of the management system; b) identification of deviations from the management system or from the procedures for performing laboratory activities; c) initiation of actions to prevent or minimize such deviations; d) reporting to laboratory management on the performance of the management system and any need for improvement; e) ensuring the effectiveness of laboratory activities. | 4.1.5 a | The laboratory shall have managerial and technical personnel who, irrespective of other responsibilities, have the authority and resources needed to carry out their duties, including the implementation, maintenance and improvement of the management system, and to identify the occurrence of departures from the management system or from the procedures for performing tests and/or calibrations, and to initiate actions to prevent or minimize such departures (see also 5.2); | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.7 | Laboratory management shall ensure that:
5.7 a | a) communication takes place regarding the effectiveness of the management system and the importance of meeting customers’ and other requirements | 4.1.5 k | The laboratory shall ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the objectives of the management system. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.7 a | a) communication takes place regarding the effectiveness of the management system and the importance of meeting customer's and other requirements | 4.1.6 | Top management shall ensure that the appropriate communication processes are established within the laboratory and that communication takes place regarding the effectiveness of the management system. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
5.7 b | b) the integrity of the management system is maintained when changes to the management system are planned and implemented. | 4.2.7 | Top management shall ensure that the integrity of the management system is maintained when changes to the management system are planned and implemented. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
6 | Resource requirements
6.1 | General The laboratory shall have available the personnel, facilities, equipment, systems and support services necessary to manage and perform its laboratory activities. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
6.2 | Personnel
OSHA A-5.2A | Technical personnel shall be: i) permanent employees or employees, who through a written contract or agreement are within the management control. ii) knowledgeable in appropriate evaluation, test procedures and test standards for the types of products covered by the lab's scope. iii) knowledgeable in the risks and hazards associated with conducting safety testing, including laboratory safety regulations, safeguards and procedures to reduce laboratory risks. | 5.2.3 |  | Accredited & Non-Accredited & In-House Cal | i) a) Confirm lab has a process in place stating this is a requirement and; i) b) review a sample copy of the contact or agreement. ii) Request records showing knowledge of a sampling of test procedures and standards within the participant's scope. A record may include past project work in testing, education and/or formal training in testing, appropriate test standards, and relevant test procedures. iii) Request records of safety training for technical staff. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.2.1 | All personnel of the laboratory, either internal or external, that could influence the laboratory activities shall act impartially, be competent and work in accordance with the laboratory’s management system. |  |  | No | Impartiality covered through agreements, competency identified 6.2.6 authorization.
6.2.2 | The laboratory shall document the competence requirements for each function influencing the results of laboratory activities, including requirements for education, qualification, training, technical knowledge, skills and experience. | 5.2.4 | The laboratory shall maintain current job descriptions for managerial, technical and key support personnel involved in tests and/or calibrations. | No | Covered by the appropriate agreement and through 6.2.5 and 6.2.6.
6.2.3 | The laboratory shall ensure that the personnel have the competence to perform laboratory activities for which they are responsible and to evaluate the significance of deviations. | 5.2.1 | The laboratory management shall ensure the competence of all who operate specific equipment, perform tests and/or calibrations, evaluate results, and sign test reports and calibration certificates. When using staff who are undergoing training, appropriate supervision shall be provided. Personnel performing specific tasks shall be qualified on the basis of appropriate education, training, experience and/or demonstrated skills, as required. | No | Covered by the appropriate agreement and through 6.2.5 and 6.2.6 (Also covers 5.5.3 from 2005) Also covered per OSHA Directive Annex A 5.2D below.
OSHA A-5.2D | The lab shall maintain records documenting competence in the particular testing, inspection, or other technical subjects, procedures, or practices that technical personnel will perform. For example, a record may include past project work in testing, education, and/or formal training in testing, appropriate test standards, and relevant test procedures |  |  | Accredited & Non-Accredited & In-House Cal | (expansion of 6.2.3 above) Sample technical staff for record of competence per the requirements. Examples of competence records can be past project work in testing, education, and/or formal training in testing, appropriate test standards, and relevant test procedures. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.2.4 | The management of the laboratory shall communicate to personnel their duties, responsibilities and authorities. |  |  | No | Covered by the appropriate agreement and through 6.2.5 and 6.2.6
6.2.5 | The laboratory shall have procedure(s) and retain records for: a) determining the competence requirements; b) selection of personnel; c) training of personnel; d) supervision of personnel; e) authorization of personnel; f) monitoring competence of personnel. | 5.2.2 | The management of the laboratory shall formulate the goals with respect to the education, training and skills of the laboratory personnel. The laboratory shall have a policy and procedures for identifying training needs and providing training of personnel. The training program shall be relevant to the present and anticipated tasks of the laboratory. The effectiveness of the training actions taken shall be evaluated. | Non-Accredited & In-House Cal | Verify procedures, training records and documentation.
OSHA A-5C | The laboratory shall i) maintain a written training program for new and current technical personnel, which shall include the proper procedures for applying new/updated test procedures and performing required tests. ii) provide current technical personnel additional training, if necessary, when test standards or procedures are updated or developed or when responsibilities have changed. iii) conduct training through appropriate training mechanisms, such as on-the-job training or formal classroom training. iv) maintain records of training for each individual who is a member of the technical staff. |  |  | Accredited & Non-Accredited & In-House Cal | i) Confirm existence of formal training programs, including one for new technical personnel ii) Ask lab to provide an example of training performed and logged due to an updated standard or test method iii) Request information on the types of training the lab uses, and proof of those types of training. iv) Sample technical staff training records to confirm this requirement. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.2.6 | The laboratory shall authorize personnel to perform specific laboratory activities, including but not limited to, the following: a) development, modification, verification and validation of methods; b) analysis of results, including statements of conformity or opinions and interpretations; c) report, review and authorization of results. | 5.2.5 | The management shall authorize specific personnel to perform particular types of sampling, test and/or calibration, to issue test reports and calibration certificates, to give opinions and interpretations and to operate particular types of equipment. The laboratory shall maintain records of the relevant authorization(s), competence, educational and professional qualifications, training, skills and experience of all technical personnel, including contracted personnel. This information shall be readily available and shall include the date on which authorization and/or competence is confirmed. | Accredited & Non-Accredited & In-House Cal | Identification of signatories & authorized personnel | Updated for OSHA Directive - now required for Accredited labs as well.
6.3 | Facilities and environmental conditions
6.3.1 | The facilities and environmental conditions shall be suitable for the laboratory activities and shall not adversely affect the validity of results. NOTE Influences that can adversely affect the validity of results can include, but are not limited to, microbial contamination, dust, electromagnetic disturbances, radiation, humidity, electrical supply, temperature, sound and vibration. | 5.3.1 | Laboratory facilities for testing and/or calibration, including but not limited to energy sources, lighting and environmental conditions, shall be such as to facilitate correct performance of the tests and/or calibrations. | Accredited & Non-Accredited & In-House Cal | If the standard in which testing is being conducted (If the test standard) identifies certain environmental conditions, those conditions must be monitored. Power quality must be assessed per DAP Client Guide Client/Customer Support Guidance-ULID-000823 (DCS:00-OP-C0036) (Power Quality), as applicable.
6.3.2 | The requirements for facilities and environmental conditions necessary for the performance of the laboratory activities shall be documented. | 5.3.1 | The technical requirements for accommodation and environmental conditions that can affect the results of tests and calibrations shall be documented. | Accredited & Non-Accredited & In-House Cal | Verify documentation and logging (monitoring) of environmental conditions.
6.3.3 | The laboratory shall monitor, control and record environmental conditions in accordance with relevant specifications, methods or procedures or where they influence the validity of the results. | 5.3.2 | The laboratory shall monitor, control and record environmental conditions as required by the relevant specifications, methods and procedures or where they influence the quality of the results. Due attention shall be paid, for example, to biological sterility, dust, electromagnetic disturbances, radiation, humidity, electrical supply, temperature, and sound and vibration levels, as appropriate to the technical activities concerned. Tests and calibrations shall be stopped when the environmental conditions jeopardize the results of the tests and/or calibrations. | Accredited & Non-Accredited & In-House Cal | Records should be available for applicable monitoring of conditions. These may include chart recordings, datasheets, data collection output, etc.
6.3.4 | Measures to control facilities shall be implemented, monitored and periodically reviewed and shall include, but not be limited to: a) access to and use of areas affecting laboratory activities; b) prevention of contamination, interference or adverse influences on laboratory activities; c) effective separation between areas with incompatible laboratory activities. | 5.3.3 | There shall be effective separation between neighboring areas in which there are incompatible activities. Measures shall be taken to prevent cross-contamination. | Accredited & Non-Accredited & In-House Cal | Auditors will determine on case-by-case basis. For example, you would not want an air ventilation system too close to a test bench when conducting a product surface temperature test.
6.3.4 | Measures to control facilities shall be implemented, monitored and periodically reviewed and shall include, but not be limited to: a) access to and use of areas affecting laboratory activities; b) prevention of contamination, interference or adverse influences on laboratory activities; c) effective separation between areas with incompatible laboratory activities. | 5.3.4 | Access to and use of areas affecting the quality of the tests and/or calibrations shall be controlled. The laboratory shall determine the extent of control based on its particular circumstances. | Accredited & Non-Accredited & In-House Cal | Auditors will verify access to test areas
6.3.5 | When the laboratory performs laboratory activities at sites or facilities outside its permanent control, it shall ensure that the requirements related to facilities and environmental conditions of this document are met. | 5.3.1 | The laboratory shall ensure that the environmental conditions do not invalidate the results or adversely affect the required quality of any measurement. Particular care shall be taken when sampling and tests and/or calibrations are undertaken at sites other than a permanent laboratory facility. | Accredited & Non-Accredited & In-House Cal | Auditors will determine if this applies based upon the scope and testing activities
6.4 | Equipment
6.4.1 | The laboratory shall have access to equipment (including, but not limited to, measuring instruments, software, measurement standards, reference materials, reference data, reagents, consumables or auxiliary apparatus) that is required for the correct performance of laboratory activities and that can influence the result. NOTE 1 A multitude of names exist for reference materials and certified reference materials, including reference standards, calibration standards, standard reference materials and quality control materials. ISO 17034 contains additional information on reference material producers (RMPs). RMPs that meet the requirements of ISO 17034 are considered to be competent. Reference materials from RMPs meeting the requirements of ISO 17034 are provided with a product information sheet/certificate that specifies, amongst other characteristics, homogeneity and stability for specified properties and, for certified reference materials, specified properties with certified values, their associated measurement uncertainty and metrological traceability. NOTE 2 ISO Guide 33 provides guidance on the selection and use of reference materials. ISO Guide 80 provides guidance to produce in-house quality control materials. | 5.5.1 | The laboratory shall be furnished with all items of sampling, measurement and test equipment required for the correct performance of the tests and/or calibrations (including sampling, preparation of test and/or calibration items, processing and analysis of test and/or calibration data). | Accredited & Non-Accredited & In-House Cal | The auditor will verify that all required equipment is available. If Equipment is borrowed and/or rented does the lab perform a verification of the calibration and functional status of the equipment prior to use?
6.4.2 | When the laboratory uses equipment outside its permanent control, it shall ensure that the requirements for equipment of this document are met. | 5.5.1 | In those cases where the laboratory needs to use equipment outside its permanent control, it shall ensure that the requirements of this handbook are met. | Accredited & Non-Accredited & In-House Cal | The auditor will verify that all required equipment is available. If Equipment is borrowed and/or rented does the lab perform a verification of the calibration and functional status of the equipment prior to use?
OSHA A-5.5A | The lab shall ensure that all equipment used for testing and evaluating products is available and in proper working order for NRTL work. However, equipment needed only occasionally for a special or unique type of product that is seldom tested, may be rented as needed. |  |  | Accredited & Non-Accredited & In-House Cal | Sample required test equipment records and confirm equipment is calibrated and in working condition. Confirm ownership of regularly used equipment. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.4.3 | The laboratory shall have a procedure for handling, transport, storage, use and planned maintenance of equipment in order to ensure proper functioning and to prevent contamination or deterioration | 5.5.6 | The laboratory shall have procedures for safe handling, transport, storage, use and planned maintenance of measuring equipment to ensure proper functioning and in order to prevent contamination or deterioration. | Accredited & Non-Accredited | A documented procedure is required.
6.4.4 | The laboratory shall verify that equipment conforms to specified requirements before being placed or returned into service | 5.5.2 | Before being placed into service, equipment (including that used for sampling) shall be calibrated or checked to establish that it meets the laboratory's specification requirements and complies with the relevant standard specifications. It shall be checked and/or calibrated before use (see 5.6). | Accredited & Non-Accredited & In-House Cal | The lab must verify their equipment, used to conduct the tests contained within their DAP Scope of Participation, meet the accuracy requirements in the product standards. If no accuracy is required or defined by the product standard(s), equipment must meet the accuracy requirements of DAP Client Guide Client/Customer Support Guidance-ULID-000821 (DCS:00-OP-C0034) (Equipment Selection). Records are required to verify this check has been made.
OSHA A-5.5B | The lab must own, lease, or rent equipment and shall have procedures requiring that new, leased, rented, and repaired equipment is calibrated prior to first use. |  |  | Accredited & Non-Accredited & In-House Cal | Confirm adequate procedure is in place requiring calibration before first use. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.4.5 | The equipment used for measurement shall be capable of achieving the measurement accuracy and/or measurement uncertainty required to provide a valid result. | 5.5.2 | Equipment and its software used for testing, calibration and sampling shall be capable of achieving the accuracy required and shall comply with specifications relevant to the tests and/or calibrations concerned. | Accredited & Non-Accredited & In-House Cal | The lab must verify their equipment, used to conduct the tests contained within their DAP Scope of Participation, meet the accuracy requirements in the product standards. If no accuracy is required or defined by the product standard(s), equipment must meet the accuracy requirements of DAP Client Guide Client/Customer Support Guidance-ULID-000821 (DCS:00-OP-C0034) (Equipment Selection). Records are required to verify this check has been made.
6.4.6 | Measuring equipment shall be calibrated when: - the measurement accuracy or measurement uncertainty affects the validity of the reported results, and/or - calibration of the equipment is required to establish the metrological traceability of the reported results. NOTE Types of equipment having an effect on the validity of the reported results can include: — those used for the direct measurement of the measurand, e.g. use of a balance to perform a mass measurement; — those used to make corrections to the measured value, e.g. temperature measurements; — those used to obtain a measurement result calculated from multiple quantities. | 5.5.2 | Before being placed into service, equipment (including that used for sampling) shall be calibrated or checked to establish that it meets the laboratory's specification requirements and complies with the relevant standard specifications. It shall be checked and/or calibrated before use (see 5.6). | Accredited & Non-Accredited & In-House Cal | The lab must verify their equipment, used to conduct the tests contained within their DAP Scope of Participation, meet the accuracy requirements in the product standards. If no accuracy is required or defined by the product standard(s), equipment must meet the accuracy requirements of DAP Client Guide Client/Customer Support Guidance-ULID-000821 (DCS:00-OP-C0034) (Equipment Selection). Records are required to verify this check has been made.
6.4.7 | The laboratory shall establish a calibration programme, which shall be reviewed and adjusted as necessary in order to maintain confidence in the status of calibration. | 5.5.2 | Calibration programs shall be established for key quantities or values of the instruments where these properties have a significant effect on the results. | Accredited & Non-Accredited & In-House Cal | Confirm existence of CAL program. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive - now required for ALL participant labs.
OSHA A-5.5C | The laboratory shall ensure that its procedures address adding, deleting, modifying or maintaining information in equipment records in an accurate and timely manner, and specify the personnel responsible for these tasks. |  |  | Accredited & Non-Accredited & In-House Cal | Confirm that there is a process in place for updating records in a timely manner and that it specifies who is responsible. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.4.8 | All equipment requiring calibration or which has a defined period of validity shall be labelled, coded or otherwise identified to allow the user of the equipment to readily identify the status of calibration or period of validity. | 5.5.8 | Whenever practicable, all equipment under the control of the laboratory and requiring calibration shall be labeled, coded or otherwise identified to indicate the status of calibration, including the date when last calibrated and the date or expiration criteria when recalibration is due. | Accredited & Non-Accredited & In-House Cal | Lab equipment labels are to contain the required information.
OSHA A-5.5D | The lab shall ensure that its procedures specify the steps for establishing calibration intervals for each type or item of equipment, and specify criteria, steps, and approvals for extending the calibration interval of an instrument. |  |  | Accredited & Non-Accredited & In-House Cal | Confirm CAL program covers these activities. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.4.9 | Equipment that has been subjected to overloading or mishandling, gives questionable results, or has been shown to be defective or outside specified requirements, shall be taken out of service. It shall be isolated to prevent its use or clearly labelled or marked as being out of service until it has been verified to perform correctly. The laboratory shall examine the effect of the defect or deviation from specified requirements and shall initiate the management of nonconforming work procedure (see 7.10). | 5.5.7 | Equipment that has been subjected to overloading or mishandling, gives suspect results, or has been shown to be defective or outside specified limits, shall be taken out of service. It shall be isolated to prevent its use or clearly labeled or marked as being out of service until it has been repaired and shown by calibration or test to perform correctly. The laboratory shall examine the effect of the defect or departure from specified limits on previous tests and/or calibrations and shall institute the "Control of nonconforming work" procedure (see 4.9). | Accredited & Non-Accredited & In-House Cal | A process should be in place to address this requirement. Records of equipment issues are needed to support compliance.
OSHA A-5.5F | The lab shall have procedures to examine the effects of defective equipment on calibration and tests. The procedures shall identify the personnel responsible for such examinations, specify their responsibilities, and provide the steps for the examination including: i) determining whether the effects are unacceptable (including the accept/reject criteria); ii) identifying the products affected; iii) analyzing the particular tests impacted for these products; and iv) determining whether retesting is required Procedures shall also specify the report or document that is prepared for this examination, the notification provided to clients and certification body when retesting is required, and the steps to follow to perform the retesting. |  |  | Accredited & Non-Accredited & In-House Cal | i) This clause is about the process surrounding defective equipment. Confirm that examination procedures are in place and that they identify the responsible personnel. ii) Confirm process is in place stating how affected products are identified. iii) Confirm process is in place for identifying which tests would be affected by the out of tolerance equipment. iv) Confirm process is in place for determining whether retesting would be required due to the specific out of tolerance issue. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
OSHA A-5.5G | If a piece of test equipment is found to be out-of-tolerance, the lab shall have procedures to: i) identify and document any product(s) tested by the out-of-tolerance equipment after the last known date the equipment was in-tolerance; ii) review and document any testing conducted using the out-of-tolerance equipment to determine if the out-of-tolerance condition impacted test results; iii) retest the products impacted by the out-of-tolerance condition and document the results; and iv) document the corrective actions taken to comply with Section 5.5.G.i, ii, and iii, above, and retain such documentation in the test equipment records and the technical files or test records for any tested products impacted by the out-of-tolerance condition. |  |  | Accredited & Non-Accredited & In-House Cal | i) This clause is all about the testing that was performed with the defective equipment. Confirm that a procedure is in place for logging affected product(s) tested. ii) Confirm that a procedure is in place to evaluate whether the out-of-tolerance condition impacted test results and if so, how trace-back occurs. iii) Confirm that a procedure is in place that requires retesting of products when testing is found to be suspect to out-of-tolerance equipment. iv) Confirm that the procedure is in place to document the actions taken as part of this process and retain the documentation. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.4.10 | When intermediate checks are necessary to maintain confidence in the performance of the equipment, these checks shall be carried out according to a procedure. | 5.5.10 | When intermediate checks are needed to maintain confidence in the calibration status of the equipment, these checks shall be carried out according to a defined procedure. | Accredited & Non-Accredited & In-House Cal | If applicable, a documented procedure is required. Equipment manuals may provide sufficient guidance to conduct an intermediate check. This requirement applies if a lab uses equipment with a one-time calibration (See Client/Customer Support Guidance-ULID-000829 (DCS:00-OP-C0045)).
6.4.11 | When calibration and reference material data include reference values or correction factors, the laboratory shall ensure the reference values and correction factors are updated and implemented, as appropriate, to meet specified requirements. | 5.5.11 | Where calibrations give rise to a set of correction factors, the laboratory shall have procedures to ensure that copies (e.g., in computer software) are correctly updated. | Accredited & Non-Accredited & In-House Cal | If applicable, a documented procedure is required. Equipment manuals may provide sufficient guidance to apply correction factors.
6.4.12 | The laboratory shall take practicable measures to prevent unintended adjustments of equipment from invalidating results. | 5.5.12 | Test and calibration equipment, including both hardware and software, shall be safeguarded from adjustments which would invalidate the test and/or calibration results. | Accredited & Non-Accredited & In-House Cal | Equipment safeguards may be tamper seals, lockouts, paint seals, etc. Software safeguards may be passwords or other means of protecting code.
6.4.13 | Records shall be retained for equipment which can influence laboratory activities. The records shall include the following, where applicable: a) the identity of equipment, including software and firmware version; b) the manufacturer’s name, type identification, and serial number or other unique identification; c) evidence of verification that equipment conforms with specified requirements; d) the current location; e) calibration dates, results of calibrations, adjustments, acceptance criteria, and the due date of the next calibration or the calibration interval; f) documentation of reference materials, results, acceptance criteria, relevant dates and the period of validity; g) the maintenance plan and maintenance carried out to date, where relevant to the performance of the equipment; h) details of any damage, malfunction, modification to, or repair of, the equipment. | 5.5.5 | Records shall be maintained of each item of equipment and its software significant to the tests and/or calibrations performed. The records shall include at least the following: a) the identity of the item of equipment and its software; b) the manufacturer's name, type identification, and serial number or other unique identification; c) checks that equipment complies with the specification (see 5.5.2); d) the current location, where appropriate; e) the manufacturer's instructions, if available, or reference to their location; f) dates, results and copies of reports and certificates of all calibrations, adjustments, acceptance criteria, and the due date of next calibration; g) the maintenance plan, where appropriate, and maintenance carried out to date; h) any damage, malfunction, modification or repair to the equipment. | Accredited & Non-Accredited & In-House Cal | Records for each piece of equipment are to be available and must contain the items required by a-h below, as applicable. An equipment record is different than a calibration record/certificate since it requires location information, accuracy confirmation, instruction manual information and maintenance and repair history. Calibration records/certificates do not typically contain this information. (also covers 5.5.3 from 2005)
OSHA A-5.5E | Equipment that does not need to be calibrated must be verified against documented specifications and/or procedures. |  |  | Accredited & Non-Accredited & In-House Cal | Sample equipment that does not require calibration (i.e. support equipment) and confirm that verification documentation is present. Note: If a lab cannot demonstrate compliance with this requirement, an Opportunity for Improvement (OFI) can be issued. After the final OSHA Directive Implementation Date of October 1, 2021, a Non-Conformance (NCR) will be written against this requirement. | Updated for OSHA Directive.
6.5 | Metrological traceability
6.5.1 | The laboratory shall establish and maintain metrological traceability of its measurement results by means of a documented unbroken chain of calibrations, each contributing to the measurement uncertainty, linking them to an appropriate reference. NOTE 1 In ISO/IEC Guide 99, metrological traceability is defined as the “property of a measurement result whereby the result can be related to a reference through a documented unbroken chain of calibrations, each contributing to the measurement uncertainty”. NOTE 2 See Annex A for additional information on metrological traceability. | 5.6.2.1.1 | For calibration laboratories, the program for calibration of equipment shall be designed and operated so as to ensure that calibrations and measurements made by the laboratory are traceable to the International System of Units (SI) (Système international d'unités). | Accredited & Non-Accredited & In-House Cal | Calibration certificates need to be in compliance with the requirements of Client/Customer Support Guidance-ULID-000819 (DCS:00-OP-C0032) (Calibration Certificate Analysis). Is the accreditation logo from the accreditation Body (e.i. A2A, PJL, ACLASS)? on the Calibration cert. Is the Accreditation Body an MRA Signatory of ILAC, APLAC or EAC, etc. Does the calibration certificate include calibration results for each measurement that will be used. For example if the time function in a dielectric tester will be used to measure time then the calibration certificate needs to include results for both voltage and time.
6.5.2 a-c | The laboratory shall ensure that measurement results are traceable to the International System of Units (SI) through: a) calibration provided by a competent laboratory; or NOTE 1 Laboratories fulfilling the requirements of this document are considered to be competent. b) certified values of certified reference materials provided by a competent producer with stated metrological traceability to the SI; or NOTE 2 Reference material producers fulfilling the requirements of ISO 17034 are considered to be competent. c) direct realization of the SI units ensured by comparison, directly or indirectly, with national or international standards. NOTE 3 Details of practical realization of the definitions of some important units are given in the SI brochure. | 5.6.2.1.1 | When using external calibration services, traceability of measurement shall be assured by the use of calibration services from laboratories that can demonstrate competence, measurement capability and traceability. | Accredited & Non-Accredited & In-House Cal | Calibration certificates need to be in compliance with the requirements of Client/Customer Support Guidance-ULID-000819 (DCS:00-OP-C0032) (Calibration Certificate Analysis). Is the accreditation logo from the accreditation Body (e.i. A2A, PJL, ACLASS)? on the Calibration cert. Is the Accreditation Body an MRA Signatory of ILAC, APLAC or EAC, etc. Does the calibration certificate include calibration results for each measurement that will be used. For example if the time function in a dielectric tester will be used to measure time then the calibration certificate needs to include results for both voltage and time.
6.5.3 | When metrological traceability to the SI units is not technically possible, the laboratory shall demonstrate metrological traceability to an appropriate reference, e.g.: a) certified values of certified reference materials provided by a competent producer; b) results of reference measurement procedures, specified methods or consensus standards that are clearly described and accepted as providing measurement results fit for their intended use and ensured by suitable comparison. | 5.6.2.1.2 | There are certain calibrations that currently cannot be strictly made in SI units. In these cases calibration shall provide confidence in measurements by establishing traceability to appropriate measurement standards such as: -- the use of certified reference materials provided by a competent supplier to give a reliable physical or chemical characterization of a material; -- the use of specified methods and/or consensus standards that are clearly described and agreed by all parties concerned. Participation in a suitable programme of interlaboratory comparisons is required where possible. | Accredited & Non-Accredited & In-House Cal | If equipment cannot be calibrated to SI Units, UL should be contacted to confirm the acceptability of calibrations to other measurement standards. UL's confirmation of acceptability needs to be kept as a record. | Updated for OSHA Directive
6.6 | Externally provided products and services
6.6.1 | The laboratory shall ensure that only suitable externally provided products and services that affect laboratory activities are used, when such products and services: NOTE Products can include, for example, measurement standards and equipment, auxiliary equipment, consumable materials and reference materials. Services can include, for example, calibration services, sampling services, testing services, facility and equipment maintenance services, proficiency testing services and assessment and auditing services. | 4.6.1 | The laboratory shall have a policy and procedure(s) for the selection and purchasing of services and supplies it uses that affect the quality of the tests and/or calibrations. | Accredited & Non-Accredited | A documented procedure required. Reference to DAP Client Guide Client/Customer Support Guidance-ULID-000820 (DCS:00-OP-C0033) (Consumables) and Client/Customer Support Guidance-ULID-000824 (DCS:00-OP-C0037) (Acceptance of Thermocouple Wire) for specific requirements.
6.6.1 a | are intended for incorporation into the laboratory’s own activities; | 4.6.1 | Procedures shall exist for the purchase, reception and storage of reagents and laboratory consumable materials relevant for the tests and calibrations. | Accredited & Non-Accredited | A documented procedure required. Reference to DAP Client Guide Client/Customer Support Guidance-ULID-000820 (DCS:00-OP-C0033) (Consumables) and Client/Customer Support Guidance-ULID-000824 (DCS:00-OP-C0037) (Acceptance of Thermocouple Wire) for specific requirements.
6.6.1 b | are provided, in part or in full, directly to the customer by the laboratory, as received from the external provider; | 4.5.1 | When a laboratory subcontracts work whether because of unforeseen reasons (e.g., workload, need for further expertise or temporary incapacity) or on a continuing basis (e.g., through permanent subcontracting, agency or franchising arrangements), this work shall be placed with a competent subcontractor. A competent subcontractor is one that, for example, complies with this handbook for the work in question. | No | DAP does not allow participants to subcontract
6.6.1 c | are used to support the operation of the laboratory. | 4.6.1 | The laboratory shall have a policy and procedure(s) for the selection and purchasing of services and supplies it uses that affect the quality of the tests and/or calibrations. Procedures shall exist for the purchase, reception and storage of reagents and laboratory consumable materials relevant for the tests and calibrations. | Accredited & Non-Accredited | A documented procedure required. Reference to DAP Client Guide Client/Customer Support Guidance-ULID-000820 (DCS:00-OP-C0033) (Consumables) and Client/Customer Support Guidance-ULID-000824 (DCS:00-OP-C0037) (Acceptance of Thermocouple Wire) for specific requirements.
6.6.2 | The laboratory shall have a procedure and retain records for: | 4.6.2 | Records of actions taken to check compliance shall be maintained. | Accredited & Non-Accredited | Reference to DAP Client Guide Client/Customer Support Guidance-ULID-000820 (DCS:00-OP-C0033) (Consumables) and Client/Customer Support Guidance-ULID-000824 (DCS:00-OP-C0037) (Acceptance of Thermocouple Wire) for specific requirements. Records of the verification of consumables are needed to support compliance. Calibration certificates must be reviewed upon receipt to confirm equipment was received by calibration house 'in-tolerance'. Also, calibration certificates will also need to be reviewed to ensure all requirements of Client Guide Client/Customer Support Guidance-ULID-000819 (DCS:00-OP-C0032) (Calibration Certificate Analysis) have been met.
6.6.2 a | defining, reviewing and approving the laboratory’s requirements for externally provided products and services; | 4.5.4 | The laboratory shall maintain a register of all subcontractors that it uses for tests and/or calibrations and a record of the evidence of compliance with this handbook for the work in question. | No | DAP does not allow participants to subcontract
6.6.2 a | defining, reviewing and approving the laboratory’s requirements for externally provided products and services; | 4.6.1 | The laboratory shall have a policy and procedure(s) for the selection and purchasing of services and supplies it uses that affect the quality of the tests and/or calibrations. | Accredited & Non-Accredited | Reference to DAP Client Guide Client/Customer Support Guidance-ULID-000820 (DCS:00-OP-C0033) (Consumables) and Client/Customer Support Guidance-ULID-000824 (DCS:00-OP-C0037) (Acceptance of Thermocouple Wire) for specific requirements. Records of the verification of consumables are needed to support compliance. Calibration certificates must be reviewed upon receipt to confirm equipment was received by calibration house 'in-tolerance'. Also, calibration certificates will also need to be reviewed to ensure all requirements of Client Guide Client/Customer Support Guidance-ULID-000819 (DCS:00-OP-C0032) (Calibration Certificate Analysis) have been met.
6.6.2 a | defining, reviewing and approving the laboratory’s requirements for externally provided products and services; | 4.6.3 | Purchasing documents for items affecting the quality of laboratory output shall contain data describing the services and supplies ordered. These purchasing documents shall be reviewed and approved for technical content prior to release. NOTE The description may include type, class, grade, precise identification, specifications, drawings, inspection instructions, other technical data including approval of test results, the quality required and the management system standard under which they were made. | No | <duplicate assessed above>
6.6.2 a | defining, reviewing and approving the laboratory’s requirements for externally provided products and services; | 4.6.4 | shall maintain records of these evaluations and list those approved. | No | <duplicate assessed above>
6.6.2 b | defining the criteria for evaluation, selection, monitoring of performance and re-evaluation of the external providers; | 4.6.4 | The laboratory shall evaluate suppliers of critical consumables, supplies and services which affect the quality of testing and calibration, and | No | Covered by client guides and agreements
6.6.2 c | ensuring that externally provided products and services conform to the laboratory's established requirements, or when applicable, to the relevant requirements of this document, before they are used or directly provided to the customer; | 4.6.2 | The laboratory shall ensure that purchased supplies and reagents and consumable materials that affect the quality of tests and/or calibrations are not used until they have been inspected or otherwise verified as complying with standard specifications or requirements defined in the methods for the tests and/or calibrations concerned. These services and supplies used shall comply with specified requirements. | No | Covered by client guides and agreements
6.6.2 d | taking any actions arising from evaluations, monitoring of performance and re-evaluations | 4.6.4 | shall maintain records of these evaluations and list those approved. | No | Covered by client guides and agreements
6.6.3 | The laboratory shall communicate its requirements to external providers for: |  |  | No | Covered by client guides and agreements
6.6.3 a | a) the products and services to be provided; | 4.6.3 | Purchasing documents for items affecting the quality of laboratory output shall contain data describing the services and supplies ordered. These purchasing documents shall be reviewed and approved for technical content prior to release. | No | Covered by client guides and agreements
6.6.3 b | b) the acceptance criteria; |  |  | No | Covered by client guides and agreements
6.6.3 c | c) competence, including any required qualification of personnel; | 4.5.1 | When a laboratory subcontracts work, whether because of unforeseen reasons (e.g. workload, need for further expertise or temporary incapacity) or on a continuing basis (e.g. through permanent subcontracting, agency or franchising arrangements), this work shall be placed with a competent subcontractor. A competent subcontractor is one that, for example, complies with this International Standard for the work in question. | No | DAP does not allow participants to subcontract
6.6.3 d | d) activities that the laboratory, or its customer, intends to perform at the external provider’s premises. |  |  | No | DAP does not allow this activity
7 | Process Requirements
7.1 | Review of requests, tenders and contracts | 4.4 | Review of requests, tenders and contracts
7.1.1 | The laboratory shall have a procedure for the review of requests, tenders and contracts. The procedure shall ensure that: | 4.4.1 | The laboratory shall establish and maintain procedures for the review of requests, tenders and contracts. The policies and procedures for these reviews leading to a contract for testing and/or calibration shall ensure that: | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.1 a | the requirements are adequately defined, documented and understood | 4.4.1 a | the requirements, including the methods to be used, are adequately defined, documented and understood (see 5.4.2); | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.1 b | the laboratory has the capability and resources to meet the requirements | 4.4.1 b | the laboratory has the capability and resources to meet the requirements; | No | DAP scopes are limited per their scope registered with UL.
7.1.1 c | where external providers are used, the requirements of 6.6 are applied and the laboratory advises the customer of the specific laboratory activities to be performed by the external provider and gains the customer’s approval; NOTE 1 It is recognized that externally provided laboratory activities can occur when: — the laboratory has the resources and competence to perform the activities, however, for unforeseen reasons is unable to undertake these in part or full; — the laboratory does not have the resources or competence to perform the activities. | 4.4.3 | The review shall also cover any work that is subcontracted by the laboratory. | No | DAP does not allow participants to subcontract
7.1.1 c | where external providers are used, the requirements of 6.6 are applied and the laboratory advises the customer of the specific laboratory activities to be performed by the external provider and gains the customer’s approval; NOTE 1 It is recognized that externally provided laboratory activities can occur when: — the laboratory has the resources and competence to perform the activities, however, for unforeseen reasons is unable to undertake these in part or full; — the laboratory does not have the resources or competence to perform the activities. | 4.5.2 | The laboratory shall advise the customer of the arrangement in writing and, when appropriate, gain the approval of the customer, preferably in writing. | No | DAP does not allow participants to subcontract
7.1.1 d | the appropriate methods or procedures are selected and are capable of meeting the customers’ requirements. | 4.4.1 c | the appropriate test and/or calibration method is selected and is capable of meeting the customers' requirements (see 5.4.2). | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.2 | The laboratory shall inform the customer when the method requested by the customer is considered to be inappropriate or out of date. | 5.4.2 | The laboratory shall inform the customer when the method proposed by the customer is considered to be inappropriate or out of date. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.3 | When the customer requests a statement of conformity to a specification or standard for the test or calibration (e.g. pass/fail, in-tolerance/out-of-tolerance), the specification or standard and the decision rule shall be clearly defined. Unless inherent in the requested specification or standard, the decision rule selected shall be communicated to, and agreed with, the customer. NOTE For further guidance on statements of conformity, see ISO/IEC Guide 98-4. |  |  | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.4 | Any differences between the request or tender and the contract shall be resolved before laboratory activities commence. Each contract shall be acceptable both to the laboratory and the customer. Deviations requested by the customer shall not impact the integrity of the laboratory or the validity of the results. | 4.4.1 | Any differences between the request or tender and the contract shall be resolved before any work commences. Each contract shall be acceptable both to the laboratory and the customer. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.5 | The customer shall be informed of any deviation from the contract | 4.4.4 | The customer shall be informed of any deviation from the contract. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.6 | If a contract is amended after work has commenced, the contract review shall be repeated and any amendments shall be communicated to all affected personnel. | 4.4.5 | If a contract needs to be amended after work has commenced, the same contract review process shall be repeated and any amendments shall be communicated to all affected personnel. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.7 | The laboratory shall cooperate with customers or their representatives in clarifying the customer’s request and in monitoring the laboratory’s performance in relation to the work performed. NOTE Such cooperation can include: | 4.7.1 | The laboratory shall be willing to cooperate with customers or their representatives in clarifying the customer’s request and in monitoring the laboratory’s performance in relation to the work performed, provided that the laboratory ensures confidentiality to other customers. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.7 a | a) providing reasonable access to relevant areas of the laboratory to witness customer-specific laboratory activities; | 4.7.1 | providing the customer or the customer's representative reasonable access to relevant areas of the laboratory for the witnessing of tests and/or calibrations performed for the customer; | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.7 b | b) preparation, packaging, and dispatch of items needed by the customer for verification purposes. | 4.7.1 b | preparation, packaging, and dispatch of test and/or calibration items needed by the customer for verification purposes. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.1.8 | Records of reviews, including any significant changes, shall be retained. Records shall also be retained of pertinent discussions with a customer relating to the customer’s requirements or the results of the laboratory activities. | 4.4.2 | Records of reviews, including any significant changes, shall be maintained. Records shall also be maintained of pertinent discussions with a customer relating to the customer's requirements or the results of the work during the period of execution of the contract. | No | CTDP - For UL purposes the test lab is only involved in the testing of its own products TPTDP – These clauses are addressed by the appropriate TPTDP agreement.
7.2 | Selection, verification and validation of methods | 5.4 | Test and calibration methods and method validation
7.2.1 | Selection and verification of methods
7.2.1.1 | The laboratory shall use appropriate methods and procedures for all laboratory activities and, where appropriate, for evaluation of the measurement uncertainty as well as statistical techniques for analysis of data. NOTE “Method” as used in this document can be considered synonymous with the term “measurement procedure” as defined in ISO/IEC Guide 99. | 5.4.1 | General The laboratory shall use appropriate methods and procedures for all tests and/or calibrations within its scope. These include sampling, handling, transport, storage and preparation of items to be tested and/or calibrated, and, where appropriate, an estimation of the measurement uncertainty as well as statistical techniques for analysis of test and/or calibration data. | In-House Cal | Required for In-House Cal. Not applicable beyond that due to the fact that UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.2.1.2 | All methods, procedures and supporting documentation, such as instructions, standards, manuals and reference data relevant to the laboratory activities, shall be kept up to date and shall be made readily available to personnel (see 8.3). | 5.4.1 | All instructions, standards, manuals and reference data relevant to the work of the laboratory shall be kept up to date and shall be made readily available to personnel (see 4.3). | In-House Cal | Labs must have access to all standards within their DAP Scope of participation. Use of UL or internally generated datasheets does not meet this requirement.
7.2.1.3 | The laboratory shall ensure that it uses the latest valid version of a method unless it is not appropriate or possible to do so. When necessary, the application of the method shall be supplemented with additional details to ensure consistent application. NOTE International, regional or national standards or other recognized specifications that contain sufficient and concise information on how to perform laboratory activities do not need to be supplemented or rewritten as internal procedures if these standards are written in a way that they can be used by the operating personnel in a laboratory. It can be necessary to provide additional documentation for optional steps in the method or additional details. | 5.4.2 | Selection of methods The laboratory shall use test and/or calibration methods, including methods for sampling, which meet the needs of the customer and which are appropriate for the tests and/or calibrations it undertakes. Methods published in international, regional or national standards shall preferably be used. The laboratory shall ensure that it uses the latest valid edition of a standard unless it is not appropriate or possible to do so. | In-House Cal | Required for In-House Cal. Not applicable beyond that due to the fact that UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.2.1.3 | The laboratory shall ensure that it uses the latest valid version of a method unless it is not appropriate or possible to do so. When necessary, the application of the method shall be supplemented with additional details to ensure consistent application. NOTE International, regional or national standards or other recognized specifications that contain sufficient and concise information on how to perform laboratory activities do not need to be supplemented or rewritten as internal procedures if these standards are written in a way that they can be used by the operating personnel in a laboratory. It can be necessary to provide additional documentation for optional steps in the method or additional details. | 5.4.2 | When necessary, the standard shall be supplemented with additional details to ensure consistent application. | No | UL provided datasheets will call out all necessary requirements
7.2.1.4 | When the customer does not specify the method to be used, the laboratory shall select an appropriate method and inform the customer of the method chosen. Methods published either in international, regional or national standards, or by reputable technical organizations, or in relevant scientific texts or journals, or as specified by the manufacturer of the equipment, are recommended. Laboratory-developed or modified methods can also be used. | 5.4.2 | When the customer does not specify the method to be used, the laboratory shall select appropriate methods that have been published either in international, regional or national standards, or by reputable technical organizations, or in relevant scientific texts or journals, or as specified by the manufacturer of the equipment. Laboratory-developed methods or methods adopted by the laboratory may also be used if they are appropriate for the intended use and if they are validated. | No | UL, as the customer, defines the test method on the datasheet. Any deviation from this method requires prior approval.
7.2.1.5 | The laboratory shall verify that it can properly perform methods before introducing them by ensuring that it can achieve the required performance. Records of the verification shall be retained. If the method is revised by the issuing body, verification shall be repeated to the extent necessary. | 5.4.2 | The laboratory shall confirm that it can properly operate standard methods before introducing the tests or calibrations. If the standard method changes, the confirmation shall be repeated. | No | Prior to the test being added to the participant's scope, it will need to be verified through technical scope expansion.
7.2.1.6 | When method development is required, this shall be a planned activity and shall be assigned to competent personnel equipped with adequate resources. As method development proceeds, periodic review shall be carried out to confirm that the needs of the customer are still being fulfilled. Any modifications to the development plan shall be approved and authorized. | 5.4.3 | Laboratory-developed methods The introduction of test and calibration methods developed by the laboratory for its own use shall be a planned activity and shall be assigned to qualified personnel equipped with adequate resources. Plans shall be updated as development proceeds and effective communication amongst all personnel involved shall be ensured. | No | Non-standard methods are not allowed without prior approval.
7.2.1.7 | Deviations from methods for all laboratory activities shall occur only if the deviation has been documented, technically justified, authorized, and accepted by the customer. | 5.4.1 | Deviation from test and calibration methods shall occur only if the deviation has been documented, technically justified, authorized, and accepted by the customer. | Accredited & Non-Accredited & In-House Cal | Non-standard methods are not allowed without prior approval.
7.2.1.7 | Deviations from methods for all laboratory activities shall occur only if the deviation has been documented, technically justified, authorized, and accepted by the customer. NOTE Customer acceptance of deviations can be agreed in advance in the contract. | 5.4.4 | Non-standard methods When it is necessary to use methods not covered by standard methods, these shall be subject to agreement with the customer and shall include a clear specification of the customer's requirements and the purpose of the test and/or calibration. The method developed shall have been validated appropriately before use. | No | Non-standard methods are not allowed without prior approval.
7.2.2 | Validation of methods | 5.4.5 | Validation of methods
7.2.2.1 | The laboratory shall validate non-standard methods, laboratory-developed methods and standard methods used outside their intended scope or otherwise modified. The validation shall be as extensive as is necessary to meet the needs of the given application or field of application. NOTE 1 Validation can include procedures for sampling, handling and transportation of test or calibration items. NOTE 2 The techniques used for method validation can be one of, or a combination of, the following: a) calibration or evaluation of bias and precision using reference standards or reference materials; b) systematic assessment of the factors influencing the result; c) testing method robustness through variation of controlled parameters, such as incubator temperature, volume dispensed; d) comparison of results achieved with other validated methods; e) interlaboratory comparisons; f) evaluation of measurement uncertainty of the results based on an understanding of the theoretical principles of the method and practical experience of the performance of the sampling or test method. | 5.4.5.2 | The laboratory shall validate non-standard methods, laboratory-designed/developed methods, standard methods used outside their intended scope, and amplifications and modifications of standard methods to confirm that the methods are fit for the intended use. The validation shall be as extensive as is necessary to meet the needs of the given application or field of application. | No | UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.2.2.2 | When changes are made to a validated method, the influence of such changes shall be determined and where they are found to affect the original validation, a new method validation shall be performed. |  |  | No | UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.2.2.3 | The performance characteristics of validated methods as assessed for the intended use, shall be relevant to the customers’ needs and consistent with specified requirements. NOTE Performance characteristics can include, but are not limited to, measurement range, accuracy, measurement uncertainty of the results, limit of detection, limit of quantification, selectivity of the method, linearity, repeatability or reproducibility, robustness against external influences or cross-sensitivity against interference from the matrix of the sample or test object, and bias. | 5.4.5.3 | The range and accuracy of the values obtainable from validated methods (e.g., the uncertainty of the results, detection limit, selectivity of the method, linearity, limit of repeatability and/or reproducibility, robustness against external influences and/or cross-sensitivity against interference from the matrix of the sample/test object), as assessed for the intended use, shall be relevant to the customers' needs. | No | UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.2.2.4 | The laboratory shall retain the following records of validation: a) the validation procedure used; b) specification of the requirements; c) determination of the performance characteristics of the method; d) results obtained; e) a statement on the validity of the method, detailing its fitness for the intended use. | 5.4.5.2 | The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use. | No | UL specifies methods on the provided datasheets and verifies the lab’s capability to perform the test through technical assessment.
7.3 | Sampling | 5.7 | Sampling
7.3.1 | The laboratory shall have a sampling plan and method when it carries out sampling of substances, materials or products for subsequent testing or calibration. The sampling method shall address the factors to be controlled to ensure the validity of subsequent testing or calibration results. The sampling plan and method shall be available at the site where sampling is undertaken. Sampling plans shall, whenever reasonable, be based on appropriate statistical methods. | 5.7.1 | The laboratory shall have a sampling plan and procedures for sampling when it carries out sampling of substances, materials or products for subsequent testing or calibration. | No | This clause is not applicable since UL, as the customer, has no requirements for sampling.
7.3.2 | The sampling method shall describe: a) the selection of samples or sites; b) the sampling plan; c) the preparation and treatment of sample(s) from a substance, material or product to yield the required item for subsequent testing or calibration. NOTE When received into the laboratory, further handling can be required as specified in 7.4. | 5.7.1 | The sampling plan as well as the sampling procedure shall be available at the location where sampling is undertaken. Sampling plans shall, whenever reasonable, be based on appropriate statistical methods. The sampling process shall address the factors to be controlled to ensure the validity of the test and calibration results. | No | This clause is not applicable since UL, as the customer, has no requirements for sampling.
7.3.3 a-g | The laboratory shall retain records of sampling data that forms part of the testing or calibration that is undertaken. These records shall include, where relevant: a) reference to the sampling method used; b) date and time of sampling; c) data to identify and describe the sample (e.g. number, amount, name); d) identification of the personnel performing sampling; e) identification of the equipment used; f) environmental or transport conditions; g) diagrams or other equivalent means to identify the sampling location, when appropriate; | 5.7.3 | The laboratory shall have procedures for recording relevant data and operations relating to sampling that forms part of the testing or calibration that is undertaken. These records shall include the sampling procedure used, the identification of the sampler, environmental conditions (if relevant) and diagrams or other equivalent means to identify the sampling location as necessary and, if appropriate, the statistics the sampling procedures are based upon. | No | This clause is not applicable since UL, as the customer, has no requirements for sampling.
7.3.3 h | h) deviations, additions to or exclusions from the sampling method and sampling plan. | 5.7.2 | Where the customer requires deviations, additions or exclusions from the documented sampling procedure, these shall be recorded in detail with the appropriate sampling data and shall be included in all documents containing test and/or calibration results, and shall be communicated to the appropriate personnel. | no | This clause is not applicable since UL, as the customer, has no requirements for sampling.
7.4 | Handling of test or calibration items | 5.8 | Handling of test or calibration items
7.4.1 | The laboratory shall have a procedure for the transportation, receipt, handling, protection, storage, retention, and disposal or return of test or calibration items, including all provisions necessary to protect the integrity of the test or calibration item, and to protect the interests of the laboratory and the customer. | 5.8.1 | The laboratory shall have procedures for the transportation, receipt, handling, protection, storage, retention and/or disposal of test and/or calibration items, including all provisions necessary to protect the integrity of the test or calibration item, and to protect the interests of the laboratory and the customer. | Accredited & Non-Accredited & In-House Cal | A documented procedure is required. | Updated for OSHA Directive - now required for Accredited labs as well.
7.4.1 | Precautions shall be taken to avoid deterioration, contamination, loss or damage to the item during handling, transporting, storing/waiting, and preparation for testing or calibration. | 5.8.4 | The laboratory shall have procedures and appropriate facilities for avoiding deterioration, loss or damage to the test or calibration item during storage, handling and preparation. | Accredited & Non-Accredited & In-House Cal | A documented procedure is required. | Updated for OSHA Directive - now required for Accredited labs as well.
7.4.1 | Handling instructions provided with the item shall be followed. | 5.8.4 | Handling instructions provided with the item shall be followed. | Accredited & Non-Accredited & In-House Cal | A documented procedure is required. | Updated for OSHA Directive - now required for Accredited labs as well.
7.4.2 | The laboratory shall have a system for the unambiguous identification of test or calibration items. The identification shall be retained while the item is under the responsibility of the laboratory. The system shall ensure that items will not be confused physically or when referred to in records or other documents. The system shall, if appropriate, accommodate a sub-division of an item or groups of items and the transfer of items. | 5.8.2 | The laboratory shall have a system for identifying test and/or calibration items. The identification shall be retained throughout the life of the item in the laboratory. The system shall be designed and operated so as to ensure that items cannot be confused physically or when referred to in records or other documents. The system shall, if appropriate, accommodate a sub-division of groups of items and the transfer of items within and from the laboratory | Accredited & Non-Accredited & In-House Cal | The method used for identification of test samples/equipment must ensure that the ID method is unique and the ID tag or other ID method is not likely to come off resulting in a mis-identification or confusion with other samples. | Updated for OSHA Directive - now required for Accredited labs as well.
7.4.3 | Upon receipt of the test or calibration item, deviations from specified conditions shall be recorded. When there is doubt about the suitability of an item for test or calibration, or when an item does not conform to the description provided, the laboratory shall consult the customer for further instructions before proceeding and shall record the results of this consultation. When the customer requires the item to be tested or calibrated acknowledging a deviation from specified conditions, the laboratory shall include a disclaimer in the report indicating which results may be affected by the deviation. | 5.8.3 | Upon receipt of the test or calibration item, abnormalities or departures from normal or specified conditions, as described in the test or calibration method, shall be recorded. When there is doubt as to the suitability of an item for test or calibration, or when an item does not conform to the description provided, or the test or calibration required is not specified in sufficient detail, the laboratory shall consult the customer for further instructions before proceeding and shall record the discussion. | Accredited & Non-Accredited & In-House Cal | Records are needed to support compliance, as applicable. | Updated for OSHA Directive - now required for Accredited labs as well.
7.4.4 | When items need to be stored or conditioned under specified environmental conditions, these conditions shall be maintained, monitored and recorded. | 5.8.4 | When items have to be stored or conditioned under specified environmental conditions, these conditions shall be maintained, monitored and recorded. | Accredited & Non-Accredited & In-House Cal | A documented procedure is required. | Updated for OSHA Directive - now required for Accredited labs as well.
7.5 | Technical Records
7.5.1 | The laboratory shall ensure that technical records for each laboratory activity contain the results, report and sufficient information to facilitate, if possible, identification of factors affecting the measurement result and its associated measurement uncertainty and enable the repetition of the laboratory activity under conditions as close as possible to the original. | 4.13.2.1 | The records for each test or calibration shall contain sufficient information to facilitate, if possible, identification of factors affecting the uncertainty and to enable the test or calibration to be repeated under conditions as close as possible to the original. | Non-Accredited | Original observations are not required to be attached to each datasheet submittal unless requested by UL, but must be kept as a record.
7.5.1 | The technical records shall include the date and the identity of personnel responsible for each laboratory activity and for checking data and results. | 4.13.2.1 | The records shall include the identity of personnel responsible for the sampling, performance of each test and/or calibration and checking of results. | Non-Accredited | Compliance is determined by review of records
7.5.1 | Original observations, data and calculations shall be recorded at the time they are made and shall be identifiable with the specific task. | 4.13.2.2 | Observations, data and calculations shall be recorded at the time they are made and shall be identifiable to the specific task. | Non-Accredited | Original data and documentation should be retained, though not required to be attached to each datasheet submittal unless requested by UL.
7.5.2 | The laboratory shall ensure that amendments to technical records can be tracked to previous versions or to original observations. Both the original and amended data and files shall be retained, including the date of alteration, an indication of the altered aspects and the personnel responsible for the alterations. | 4.13.2.1 | The laboratory shall retain records of original observations, derived data and sufficient information to establish an audit trail, calibration records, staff records and a copy of each test report or calibration certificate issued, for a defined period. | Non-Accredited | Compliance is determined by review of amended records
7.6 | Evaluation of measurement uncertainty | 5.4.6 | Estimation of uncertainty of measurement
7.6.1 | Laboratories shall identify the contributions to measurement uncertainty. When evaluating measurement uncertainty, all contributions that are of significance, including those arising from sampling, shall be taken into account using appropriate methods of analysis. | 5.4.6.3 | When estimating the uncertainty of measurement, all uncertainty components which are of importance in the given situation shall be taken into account using appropriate methods of analysis. | No | This clause is not applicable as measurement uncertainty is not required for data submitted in support of UL LLC certification decisions unless defined in the standard.
7.6.2 | A laboratory performing calibrations, including of its own equipment, shall evaluate the measurement uncertainty for all calibrations. | 5.4.6.1 | A calibration laboratory, or a testing laboratory performing its own calibrations, shall have and shall apply a procedure to estimate the uncertainty of measurement for all calibrations and types of calibrations. | No | This clause is not applicable as measurement uncertainty is not required for data submitted in support of UL LLC certification decisions unless defined in the standard.
7.6.3 | A laboratory performing testing shall evaluate measurement uncertainty. Where the test method precludes rigorous evaluation of measurement uncertainty, an estimation shall be made based on an understanding of the theoretical principles or practical experience of the performance of the method. NOTE 1 In those cases where a well-recognized test method specifies limits to the values of the major sources of measurement uncertainty and specifies the form of presentation of the calculated results, the laboratory is considered to have satisfied 7.6.3 by following the test method and reporting instructions. NOTE 2 For a particular method where the measurement uncertainty of the results has been established and verified, there is no need to evaluate measurement uncertainty for each result if the laboratory can demonstrate that the identified critical influencing factors are under control. NOTE 3 For further information, see ISO/IEC Guide 98-3, ISO 21748 and the ISO 5725 series. | 5.4.6.2 | Testing laboratories shall have and shall apply procedures for estimating uncertainty of measurement. In certain cases the nature of the test method may preclude rigorous, metrologically and statistically valid, calculation of uncertainty of measurement. In these cases the laboratory shall at least attempt to identify all the components of uncertainty and make a reasonable estimation, and shall ensure that the form of reporting of the result does not give a wrong impression of the uncertainty. Reasonable estimation shall be based on knowledge of the performance of the method and on the measurement scope and shall make use of, for example, previous experience and validation data. | In-House Cal | This clause is not applicable unless in-house Calibration is performed. Measurement uncertainty is not required for data submitted in support of UL LLC certification decisions unless defined in the standard.
7.7 | Ensuring the validity of results
7.7.1 | The laboratory shall have a procedure for monitoring the validity of results. The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to review the results. This monitoring shall be planned and reviewed and shall include, where appropriate, but not be limited to: | 5.9.1 | The laboratory shall have quality control procedures for monitoring the validity of tests and calibrations undertaken. The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to the reviewing of the results. This monitoring shall be planned and reviewed and may include, but not be limited to, the following: NOTE The selected methods should be appropriate for the type and volume of the work undertaken. | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 a | a) use of reference materials or quality control materials; | 5.9.1 a | a) regular use of certified reference materials and/or internal quality control using secondary reference materials; | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 b | b) use of alternative instrumentation that has been calibrated to provide traceable results; |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 c | c) functional check(s) of measuring and testing equipment; |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 d | d) use of check or working standards with control charts, where applicable; |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 e | e) intermediate checks on measuring equipment; |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 f | f) replicate tests or calibrations using the same or different methods; | 5.9.1 c | c) replicate tests or calibrations using the same or different methods; | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 g | g) retesting or recalibration of retained items; | 5.9.1 d | d) retesting or recalibration of retained items; | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 h | h) correlation of results for different characteristics of an item; | 5.9.1 e | e) correlation of results for different characteristics of an item. | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 i | i) review of reported results; |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 j | j) intralaboratory comparisons | 5.9.1 b | b) participation in interlaboratory comparison or proficiency-testing programs; | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.1 k | k) testing of blind sample(s). |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.2 | The laboratory shall monitor its performance by comparison with results of other laboratories, where available and appropriate. This monitoring shall be planned and reviewed and shall include, but not be limited to, either or both of the following: a) participation in proficiency testing; NOTE ISO/IEC 17043 contains additional information on proficiency tests and proficiency testing providers. Proficiency testing providers that meet the requirements of ISO/IEC 17043 are considered to be competent. b) participation in interlaboratory comparisons other than proficiency testing. |  |  | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.7.3 | Data from monitoring activities shall be analyzed, used to control and, if applicable, improve the laboratory’s activities. If the results of the analysis of data from monitoring activities are found to be outside pre-defined criteria, appropriate action shall be taken to prevent incorrect results from being reported. | 5.9.2 | Quality control data shall be analyzed and, where they are found to be outside pre-defined criteria, planned action shall be taken to correct the problem and to prevent incorrect results from being reported. | No | This clause is not applicable due to the fact that UL requires that the participant’s technical scope is assessed on a regular cycle to ensure the quality of the test data received and reserves the right to retest products at any time.
7.8 | Reporting of results | 5.1 | Reporting of results
7.8.1 | General | 5.10.1 | General
7.8.1.(1-2) | The results shall be reviewed and authorized prior to release. The results shall be provided accurately, clearly, unambiguously and objectively, usually in a report (e.g. a test report or a calibration certificate or report of sampling), and shall include all the information agreed with the customer and necessary for the interpretation of the results and all information required by the method used. All issued reports shall be retained as technical records. NOTE 1 For the purposes of this document, test reports and calibration certificates are sometimes referred to as test certificates and calibration reports, respectively. NOTE 2 Reports can be issued as hard copies or by electronic means, provided that the requirements of this document are met. | 5.10.1 | The results of each test, calibration, or series of tests or calibrations carried out by the laboratory shall be reported accurately, clearly, unambiguously and objectively, and in accordance with any specific instructions in the test or calibration methods. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.1.(1-2) | The results shall be reviewed and authorized prior to release. The results shall be provided accurately, clearly, unambiguously and objectively, usually in a report (e.g. a test report or a calibration certificate or report of sampling), and shall include all the information agreed with the customer and necessary for the interpretation of the results and all information required by the method used. All issued reports shall be retained as technical records. NOTE 1 For the purposes of this document, test reports and calibration certificates are sometimes referred to as test certificates and calibration reports, respectively. NOTE 2 Reports can be issued as hard copies or by electronic means, provided that the requirements of this document are met. | 5.10.1 | The results shall be reported, usually in a test report or a calibration certificate (see Note 1), and shall include all the information requested by the customer and necessary for the interpretation of the test or calibration results and all information required by the method used. This information is normally that required by 5.10.2, and 5.10.3 or 5.10.4. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.1.(1-2) | The results shall be reviewed and authorized prior to release. The results shall be provided accurately, clearly, unambiguously and objectively, usually in a report (e.g. a test report or a calibration certificate or report of sampling), and shall include all the information agreed with the customer and necessary for the interpretation of the results and all information required by the method used. All issued reports shall be retained as technical records. NOTE 1 For the purposes of this document, test reports and calibration certificates are sometimes referred to as test certificates and calibration reports, respectively. NOTE 2 Reports can be issued as hard copies or by electronic means, provided that the requirements of this document are met. | 5.10.7 | Electronic transmission of results In the case of transmission of test or calibration results by telephone, telex, facsimile or other electronic or electromagnetic means, the requirements of this International Standard shall be met (see also 5.4.7). | No | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.1.3 | When agreed with the customer, the results may be reported in a simplified way. Any information listed in 7.8.2 to 7.8.7 that is not reported to the customer shall be readily available. | 5.10.1 | In the case of tests or calibrations performed for internal customers, or in the case of a written agreement with the customer, the results may be reported in a simplified way. Any information listed in 5.10.2 to 5.10.4 which is not reported to the customer shall be readily available in the laboratory which carried out the tests and/or calibrations. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2 | Common requirements for reports (test, calibration or sampling) | 5.10.2 | Test reports and calibration certificates
7.8.2.1 | Each report shall include at least the following information, unless the laboratory has valid reasons for not doing so, thereby minimizing any possibility of misunderstanding or misuse: NOTE Including a statement specifying that the report shall not be reproduced except in full without approval of the laboratory can provide assurance that parts of a report are not taken out of context. | 5.10.2 | Test reports and calibration certificates Each test report or calibration certificate shall include at least the following information, unless the laboratory has valid reasons for not doing so: | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 a | a) a title (e.g. “Test Report”, “Calibration Certificate” or “Report of Sampling”); | 5.10.2 a | a) a title (e.g., "Test Report" or "Calibration Certificate"); | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 b | b) the name and address of the laboratory; | 5.10.2 b | b) the name and address of the laboratory, | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 c | c) the location of performance of the laboratory activities, including when performed at a customer facility or at sites away from the laboratory’s permanent facilities, or in associated temporary or mobile facilities; | 5.10.2 b | b) and the location where the tests and/or calibrations were carried out, if different from the address of the laboratory; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 d | d) unique identification that all its components are recognized as a portion of a complete report and a clear identification of the end; | 5.10.2 c | c) unique identification of the test report or calibration certificate (such as the serial number), and on each page an identification in order to ensure that the page is recognized as a part of the test report or calibration certificate, and a clear identification of the end of the test report or calibration certificate; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 e | e) the name and contact information of the customer; | 5.10.2 d | d) the name and address of the customer; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 f | f) identification of the method used; | 5.10.2 e | e) identification of the method used; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 g | g) a description, unambiguous identification, and, when necessary, the condition of the item; | 5.10.2 f | f) a description of, the condition of, and unambiguous identification of the item(s) tested or calibrated; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 h | h) the date of receipt of the test or calibration item(s), and the date of sampling, where this is critical to the validity and application of the results; | 5.10.2 g | g) the date of receipt of the test or calibration item(s) where this is critical to the validity and application of the results, and the date(s) of performance of the test or calibration; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 i | i) the date(s) of performance of the laboratory activity; | 5.10.2 g | g) the date of receipt of the test or calibration item(s) where this is critical to the validity and application of the results, and the date(s) of performance of the test or calibration; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 j | j) the date of issue of the report; |  |  | Accredited & Non-Accredited & In-House Cal | Date of issue shall be the signatory date.
7.8.2.1 k | k) reference to the sampling plan and sampling method used by the laboratory or other bodies where these are relevant to the validity or application of the results; | 5.10.2 h | h) reference to the sampling plan and procedures used by the laboratory or other bodies where these are relevant to the validity or application of the results; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 l | l) a statement to the effect that the results relate only to the items tested, calibrated or sampled; | 5.10.2 k | k) where relevant, a statement to the effect that the results relate only to the items tested or calibrated. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 l | l) a statement to the effect that the results relate only to the items tested, calibrated or sampled; | 5.10.4.2 | The calibration certificate shall relate only to quantities and the results of functional tests. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 m | m) the results with, where appropriate, the units of measurement; | 5.10.2 i | i) the test or calibration results with, where appropriate, the units of measurement; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 n | n) additions to, deviations, or exclusions from the method; | 5.10.3.1 a | a) deviations from, additions to, or exclusions from the test method, and information on specific test conditions, such as environmental conditions; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 o | o) identification of the person(s) authorizing the report; | 5.10.2 j | j) the name(s), function(s) and signature(s) or equivalent identification of person(s) authorizing the test report or calibration certificate; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.2.1 p | p) clear identification when results are from external providers. | 5.10.6 | Testing and calibration results obtained from subcontractors When the test report contains results of tests performed by subcontractors, these results shall be clearly identified. The subcontractor shall report the results in writing or electronically. When a calibration has been subcontracted, the laboratory performing the work shall issue the calibration certificate to the contracting laboratory. | No | Subcontracting is not allowed in the Data Acceptance Program.
7.8.2.2 | The laboratory shall be responsible for all the information provided in the report, except when information is provided by the customer. Data provided by a customer shall be clearly identified. In addition, a disclaimer shall be put on the report when the information is supplied by the customer and can affect the validity of results. Where the laboratory has not been responsible for the sampling stage (e.g. the sample has been provided by the customer), it shall state in the report that the results apply to the sample as received. |  |  | Accredited & Non-Accredited & In-House Cal | TPTDP/PPP only - Confirm that customer provided information is identified.
7.8.3 | Specific requirements for test reports | 5.10.3 | Test reports
7.8.3.1 | In addition to the requirements listed in 7.8.2, test reports shall, where necessary for the interpretation of the test results, include the following: | 5.10.3.1 | In addition to the requirements listed in 5.10.2, test reports shall, where necessary for the interpretation of the test results, include the following: | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.1 a | a) information on specific test conditions, such as environmental conditions; | 5.10.3.1 a | a) deviations from, additions to, or exclusions from the test method, and information on specific test conditions, such as environmental conditions; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.1 b | b) where relevant, a statement of conformity with requirements or specifications (see 7.8.6); | 5.10.3.1 b | b) where relevant, a statement of compliance/non-compliance with requirements and/or specifications; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.1 c | c) where applicable, the measurement uncertainty presented in the same unit as that of the measurand or in a term relative to the measurand (e.g. percent) when: — it is relevant to the validity or application of the test results; — a customer’s instruction so requires, or — the measurement uncertainty affects conformity to a specification limit; | 5.10.3.1 c | c) where applicable, a statement on the estimated uncertainty of measurement; information on uncertainty is needed in test reports when it is relevant to the validity or application of the test results, when a customer's instruction so requires, or when the uncertainty affects compliance to a specification limit; | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.1 d | d) where appropriate and needed, opinions and interpretations (see 5.10.5); | 5.10.3.1 d | d) where appropriate and needed, opinions and interpretations (see 5.10.5); | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.1 e | e) additional information that may be required by specific methods, customers or groups of customers. | 5.10.3.1 e | e) additional information which may be required by specific methods, customers or groups of customers. | Accredited & Non-Accredited & In-House Cal | Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000818 (DCS:00-OP-C0025) (Data Recording and Reporting Requirements) for guidance on datasheet requirements. Please refer to DAP Client Guide Client/Customer Support Guidance-ULID-000825 (DCS:00-OP-C0038) (In-house Calibration) for guidance on in-house calibration records.
7.8.3.2 | Where the laboratory is responsible for the sampling activity, test reports shall meet the requirements listed in 7.8.5 where necessary for the interpretation of test results. |  |  | No | We do not perform sampling as part of DAP
7.8.4 | Specific requirements for calibration certificates
7.8.4.1 | In addition to the requirements listed in 7.8.2, calibration certificates shall include the following: | 5.10.4.1 | In addition to the requirements listed in 5.10.2, calibration certificates shall include the following, where necessary for the interpretation of calibration results: | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 a | a) the measurement uncertainty of the measurement result presented in the same unit as that of the measurand or in a term relative to the measurand (e.g. percent); | 5.10.4.1 b | b) the uncertainty of measurement and/or a statement of compliance with an identified metrological specification or clauses thereof; | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 b | b) the conditions (e.g. environmental) under which the calibrations were made that have an influence on the measurement results; | 5.10.4.1 a | a) the conditions (e.g., environmental) under which the calibrations were made that have an influence on the measurement results; | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 c | c) a statement identifying how the measurements are metrologically traceable (see Annex A); | 5.10.4.1 c | c) evidence that the measurements are traceable (see Note 2 in 5.6.2.1.1). | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 d | d) the results before and after any adjustment or repair, if available; | 5.10.4.3 | When an instrument for calibration has been adjusted or repaired, the calibration results before and after adjustment or repair, if available, shall be reported. | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 e | e) where relevant, a statement of conformity with requirements or specifications (see 7.8.6); | 5.10.4.2 | If a statement of compliance with a specification is made, this shall identify which clauses of the specification are met or not met. | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.1 f | f) where appropriate, opinions and interpretations (see 7.8.7). |  |  | In-House Cal | Other than In-House Cal, these clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.4.2 | Where the laboratory is responsible for the sampling activity, calibration certificates shall meet the requirements listed in 7.8.5 where necessary for the interpretation of calibration results. |  |  | No | UL, as the customer has no requirements for sampling.
7.8.4.3 | A calibration certificate or calibration label shall not contain any recommendation on the calibration interval, except where this has been agreed with the customer. | 5.10.4.4 | A calibration certificate (or calibration label) shall not contain any recommendation on the calibration interval except where this has been agreed with the customer. This requirement may be superseded by legal regulations. | No | Not applicable as DAP does not cover the act of calibration.
7.8.5 | Reporting sampling – specific requirements Where the laboratory is responsible for the sampling activity, in addition to the requirements listed in 7.8.2, reports shall include the following, where necessary for the interpretation of results: a) the date of sampling; b) unique identification of the item or material sampled (including the name of the manufacturer, the model or type of designation and serial numbers, as appropriate); c) the location of sampling, including any diagrams, sketches or photographs; d) a reference to the sampling plan and sampling method; e) details of any environmental conditions during sampling that affect the interpretation of the results; f) information required to evaluate measurement uncertainty for subsequent testing or calibration. | 5.10.3.2 | In addition to the requirements listed in 5.10.2 and 5.10.3.1, test reports containing the results of sampling shall include the following, where necessary for the interpretation of test results: a) the date of sampling; b) unambiguous identification of the substance, material or product sampled (including the name of the manufacturer, the model or type of designation and serial numbers as appropriate); c) the location of sampling, including any diagrams, sketches or photographs; d) a reference to the sampling plan and procedures used; e) details of any environmental conditions during sampling that may affect the interpretation of the test results; f) any standard or other specification for the sampling method or procedure, and deviations, additions to or exclusions from the specification concerned. | No | UL, as the customer has no requirements for sampling.
7.8.6 | Reporting statements of conformity
7.8.6.1 | When a statement of conformity to a specification or standard is provided, the laboratory shall document the decision rule employed, taking into account the level of risk (such as false accept and false reject and statistical assumptions) associated with the decision rule employed, and apply the decision rule. NOTE Where the decision rule is prescribed by the customer, regulations or normative documents, a further consideration of the level of risk is not necessary. | 5.10.4.2 | When statements of compliance are made, the uncertainty of measurement shall be taken into account. | No | UL does not request or allow customers to make statements of conformity.
7.8.6.2 | The laboratory shall report on the statement of conformity, such that the statement clearly identifies: a) to which results the statement of conformity applies; b) which specifications, standards or parts thereof are met or not met; c) the decision rule applied(unless it is inherent in the requested specification or standard). NOTE For further information, see ISO/IEC Guide 98-4. |  |  | No | UL does not request or allow customers to make statements of conformity.
7.8.7 | Reporting opinions and interpretations
7.8.7.1 | When opinions and interpretations are expressed, the laboratory shall ensure that only personnel authorized for the expression of opinions and interpretations release the respective statement. The laboratory shall document the basis upon which the opinions and interpretations have been made. NOTE It is important to distinguish opinions and interpretations from statements of inspections and product certifications as intended in ISO/IEC 17020 and ISO/IEC 17065, and from statements of conformity as referred to in 7.8.6. | 5.10.5 | When opinions and interpretations are included, the laboratory shall document the basis upon which the opinions and interpretations have been made. Opinions and interpretations shall be clearly marked as such in a test report. NOTE 1 Opinions and interpretations should not be confused with inspections and product certifications as intended in ISO/IEC 17020 and ISO/IEC Guide 65. | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.7.2 | The opinions and interpretations expressed in reports shall be based on the results obtained from the tested or calibrated item and shall be clearly identified as such. | 5.10.5 | NOTE 2 Opinions and interpretations included in a test report may comprise, but not be limited to, the following: --an opinion on the statement of compliance/noncompliance of the results with requirements; -- fulfilment of contractual requirements; -- recommendations on how to use the results; -- guidance to be used for improvements. | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.7.3 | When opinions and interpretations are directly communicated by dialogue with the customer, a record of the dialogue shall be retained. | 5.10.5 | NOTE 3 In many cases it might be appropriate to communicate the opinions and interpretations by direct dialogue with the customer. Such dialogue should be written down. | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.8 | Amendments to reports
7.8.8.1 | When an issued report needs to be changed, amended or re-issued, any change of information shall be clearly identified and, where appropriate, the reason for the change included in the report. |  |  | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.8.2 | Amendments to a report after issue shall be made only in the form of a further document, or data transfer, which includes the statement “Amendment to Report, serial number... [or as otherwise identified]”, or an equivalent form of wording. Such amendments shall meet all the requirements of this document. | 5.10.9 | Material amendments to a test report or calibration certificate after issue shall be made only in the form of a further document, or data transfer, which includes the statement: “Supplement to Test Report [or Calibration Certificate], serial number... [or as otherwise identified]”, or an equivalent form of wording. | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.8.8.3 | When it is necessary to issue a complete new report, this shall be uniquely identified and shall contain a reference to the original that it replaces. | 5.10.9 | When it is necessary to issue a complete new test report or calibration certificate, this shall be uniquely identified and shall contain a reference to the original that it replaces. | No | These clauses are not applicable since UL, as the customer of the lab, specifies the form and process for submitting the test data to UL.
7.9 | Complaints
7.9.1 | The laboratory shall have a documented process to receive, evaluate and make decisions on complaints. | 4.8 | The laboratory shall have a policy and procedure for the resolution of complaints received from customers or other parties. | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.2 | A description of the handling process for complaints shall be available to any interested party on request. Upon receipt of a complaint, the laboratory shall confirm whether the complaint relates to laboratory activities that it is responsible for and, if so, shall deal with it. The laboratory shall be responsible for all decisions at all levels of the handling process for complaints. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.3 | The process for handling complaints shall include at least the following elements and methods: |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.3 a | description of the process for receiving, validating, investigating the complaint, and deciding what actions are to be taken in response to it; |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.3 b | tracking and recording complaints, including actions undertaken to resolve them; | 4.8 | Records shall be maintained of all complaints and of the investigations and corrective actions taken by the laboratory (see also 4.11). | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.3 c | ensuring that any appropriate action is taken. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.4 | The laboratory receiving the complaint shall be responsible for gathering and verifying all necessary information to validate the complaint. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.5 | Whenever possible, the laboratory shall acknowledge receipt of the complaint, and provide the complainant with progress reports and the outcome. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.6 | The outcomes to be communicated to the complainant shall be made by, or reviewed and approved by, individual(s) not involved in the original laboratory activities in question. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.9.7 | Whenever possible, the laboratory shall give formal notice of the end of the complaint handling to the complainant. |  |  | No | CTDP - When manufacturer generated test data is submitted for UL purposes, manufacturer test labs have no other customers other than UL. CTDP & TPTDP - UL does not “complain” to clients since this is viewed as being passive. Instead, UL is more direct and maintains its own records identifying any noncompliances with the laboratory and requires a resolution acceptable to UL using the manufacturer’s corrective action process which is assessed annually.
7.10 | Nonconforming Work
7.10.1 | The laboratory shall have a procedure that shall be implemented when any aspect of its laboratory activities or results of this work do not conform to its own procedures or the agreed requirements of the customer (e.g. equipment or environmental conditions are out of specified limits, results of monitoring fail to meet specified criteria). The procedure shall ensure that: | 4.9.1 | The laboratory shall have a policy and procedures that shall be implemented when any aspect of its testing and/or calibration work, or the results of this work, do not conform to its own procedures or the agreed requirements of the customer. The policy and procedures shall ensure that: | Accredited & Non-Accredited & In-House Cal | A procedure must be in place that addresses 7.10.1.a-f
7.10.1 a | a) responsibilities and authorities for the management of nonconforming work are defined; | 4.9.1 a | a) the responsibilities and authorities for the management of nonconforming work are designated and actions (including halting of work and withholding of test reports and calibration certificates, as necessary) are defined and taken when nonconforming work is identified; | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.1 b | b) actions (including halting or repeating of work and withholding of reports, as necessary) are based upon the risk levels established by the laboratory; | 4.9.1 c | c) correction is taken immediately, together with any decision about the acceptability of the nonconforming work; | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.1 c | c) evaluation is made of the significance of the nonconforming work, including an impact analysis on previous results; | 4.9.1 b | b) an evaluation of the significance of the nonconforming work is made; | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.1 d | d) a decision is taken on the acceptability of the nonconforming work; | 4.9.1 c | c) correction is taken immediately, together with any decision about the acceptability of the nonconforming work; | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.1 e | e) where necessary, the customer is notified and work is recalled; | 4.9.1 d | d) where necessary, the customer is notified and work is recalled; | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.1 f | f) the responsibility for authorizing the resumption of work is defined. | 4.9.1 e | e) the responsibility for authorizing the resumption of work is defined. | Accredited & Non-Accredited & In-House Cal | Confirm existence of the procedure and its ability to address the items in a-f
7.10.2 | The laboratory shall retain records of nonconforming work and actions as specified in 7.10.1, bullets b) to f). |  |  | Accredited & Non-Accredited & In-House Cal | Confirm existence of a log of non-conforming work.
7.10.3 | Where the evaluation indicates that the nonconforming work could recur, or that there is doubt about the conformity of the laboratory’s operations with its own management system, the laboratory shall implement corrective action. | 4.9.2 | Where the evaluation indicates that the nonconforming work could recur or that there is doubt about the compliance of the laboratory's operations with its own policies and procedures, the corrective action procedures given in 4.11 shall be promptly followed. | Accredited & Non-Accredited & In-House Cal | The Control of Nonconforming Testing Procedure must reference the Corrective Action Procedure and in what circumstances it should be followed.
7.11 | Control of data and information management | 5.4.7
7.11.1 | The laboratory shall have access to the data and information needed to perform laboratory activities. |  |  | Accredited & Non-Accredited & In-House Cal | Confirm access to Standards, SOPs, User Manuals, Datasheets, Work Instructions, etc.
7.11.2 | The laboratory information management system(s) used for the collection, processing, recording, reporting, storage or retrieval of data shall be validated for functionality, including the proper functioning of interfaces within the laboratory information management system(s) by the laboratory before introduction. Whenever there are any changes, including laboratory software configuration or modifications to commercial off-the-shelf software, they shall be authorized, documented and validated before implementation. NOTE 1 In this document “laboratory information management system(s)” includes the management of data and information contained in both computerized and non-computerized systems. Some of the requirements can be more applicable to computerized systems than to non-computerized systems. NOTE 2 Commercial off-the-shelf software in general use within its designed application range can be considered to be sufficiently validated. | 5.4.7.2 a | When computers or automated equipment are used for the acquisition, processing, recording, reporting, storage or retrieval of test or calibration data, the laboratory shall ensure that: a) computer software developed by the user is documented in sufficient detail and is suitably validated as being adequate for use; NOTE Commercial off-the-shelf software (e.g. word processing, database and statistical programmes) in general use within their designed application range may be considered to be sufficiently validated. However, laboratory software configuration/modifications should be validated as in 5.4.7.2 a). | Accredited & Non-Accredited & In-House Cal | Lab should be able to show proof of validation appropriate for the software used for lab activities, and retain records of this validation.
7.11.3 a-b,d-e | The laboratory information management system(s) shall: a) be protected from unauthorized access; b) be safeguarded against tampering and loss; d) be maintained in a manner that ensures the integrity of the data and information; e) include recording system failures and the appropriate immediate and corrective actions. | 5.4.7.2 b | b) procedures are established and implemented for protecting the data; such procedures shall include, but not be limited to, integrity and confidentiality of data entry or collection, data storage, data transmission and data processing; | Accredited & Non-Accredited & In-House Cal | Confirm that safeguards are in place.
7.11.3 c | c) The laboratory information management system(s) shall be operated in an environment that complies with provider or laboratory specifications or, in the case of non-computerized systems, provides conditions which safeguard the accuracy of manual recording and transcription | 5.4.7.2 c | c) computers and automated equipment are maintained to ensure proper functioning and are provided with the environmental and operating conditions necessary to maintain the integrity of test and calibration data. | Accredited & Non-Accredited & In-House Cal | If applicable, a process should exist for calculation and data transfers checks, and records available.
7.11.4 | When a laboratory information management system is managed and maintained off-site or through an external provider, the laboratory shall ensure that the provider or operator of the system complies with all applicable requirements of this document. |  |  | Accredited & Non-Accredited & In-House Cal | If applicable, confirm that external provider has safeguards and appropriate processes in place.
7.11.5 | The laboratory shall ensure that instructions, manuals and reference data relevant to the laboratory information management system(s) are made readily available to personnel. |  |  | Accredited & Non-Accredited & In-House Cal | Staff shall know how to access and use the LIMS and adequate instructions are available.
7.11.6 | Calculations and data transfers shall be checked in an appropriate and systematic manner. | 5.4.7.1 | Calculations and data transfers shall be subject to appropriate checks in a systematic manner. | Accredited & Non-Accredited & In-House Cal | A process should exist for calculation and data transfers checks, and records available.
8 | Management system requirements
8.1 | Options
8.1.1 | General The laboratory shall establish, document, implement and maintain a management system that is capable of supporting and demonstrating the consistent achievement of the requirements of this document and assuring the quality of the laboratory results. In addition to meeting the requirements of Clauses 4 to 7, the laboratory shall implement a management system in accordance with Option A or Option B. NOTE See Annex B for more information. | 4.2.1 | The laboratory shall establish, implement and maintain a management system appropriate to the scope of its activities. | No | Overview Clause
8.1.2 | Option A As a minimum, the management system of the laboratory shall address the following: — management system documentation (see 8.2); — control of management system documents (see 8.3); — control of records (see 8.4); — actions to address risks and opportunities (see 8.5); — improvement (see 8.6); — corrective actions (see 8.7); — internal audits (see 8.8); — management reviews (see 8.9). |  |  | No | Overview Clause
8.1.3 | Option B A laboratory that has established and maintains a management system, in accordance with the requirements of ISO 9001, and that is capable of supporting and demonstrating the consistent fulfilment of the requirements of Clauses 4 to 7, also fulfils at least the intent of the management system requirements specified in 8.2 to 8.9. |  |  | No | Overview Clause
8.2 | Management system documentation (Option A)
8.2.1 | Laboratory management shall establish, document, and maintain policies and objectives for the fulfilment of the purposes of this document and shall ensure that the policies and objectives are acknowledged and implemented at all levels of the laboratory organization. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.2.2 | The policies and objectives shall address the competence, impartiality and consistent operation of the laboratory. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.2.3 | Laboratory management shall provide evidence of commitment to the development and implementation of the management system and to continually improving its effectiveness. | 4.2.3 | Top management shall provide evidence of commitment to the development and implementation of the management system and to continually improve its effectiveness. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.2.4 | All documentation, processes, systems, records, related to the fulfilment of the requirements of this document shall be included in, referenced from, or linked to the management system. | 4.2.5 | The quality manual shall include or make reference to the supporting procedures including technical procedures. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.2.5 | All personnel involved in laboratory activities shall have access to the parts of the management system documentation and related information that are applicable to their responsibilities. | 4.2.1 | The system’s documentation shall be communicated to, understood by, available to, and implemented by the appropriate personnel. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.3 | Control of management system documents (Option A)
8.3.1 | The laboratory shall control the documents (internal and external) that relate to the fulfilment of this document. NOTE In this context, “documents” can be policy statements, procedures, specifications, manufacturer’s instructions, calibration tables, charts, text books, posters, notices, memoranda, drawings, plans, etc. These can be on various media, such as hard copy or digital. |  |  | No | UL provides the necessary documents to be used in completion of testing.
8.3.2 | The laboratory shall ensure that:
8.3.2 a | documents are approved for adequacy prior to issue by authorized personnel. | 4.3.2.1 | All documents issued to personnel in the laboratory as part of the management system shall be reviewed and approved for use by authorized personnel prior to issue. | Non-Accredited | Records of reviews and approvals are needed to support compliance.
8.3.2 b | documents are periodically reviewed, and updated as necessary; | 4.3.2.2 b | The procedure(s) adopted shall ensure documents are periodically reviewed and, where necessary, revised to ensure continuing suitability and compliance with applicable requirements; | Non-Accredited | There shall be a process that defines the period at which documents are reviewed.
8.3.2 c | changes and the current revision status of documents are identified; | 4.3.2.1 | A master list or an equivalent document control procedure identifying the current revision status and distribution of documents in the management system shall be established and be readily available to preclude the use of invalid and/or obsolete documents. | Non-Accredited | Documents shall include revision numbers and change summaries.
8.3.2 c | changes and the current revision status of documents are identified; | 4.3.3.2 | Where practicable, the altered or new text shall be identified in the document or the appropriate attachments. | Non-Accredited | Documents shall include revision numbers and change summaries.
8.3.2 c | changes and the current revision status of documents are identified; | 4.3.3.3 | If the laboratory's document control system allows for the amendment of documents by hand pending the re-issue of the documents, the procedures and authorities for such amendments shall be defined. Amendments shall be clearly marked, initialed and dated. A revised document shall be formally re-issued as soon as practicable. | Non-Accredited | Documents shall include revision numbers and change summaries.
8.3.2 d | d) relevant versions of applicable documents are available at points of use and, where necessary, their distribution is controlled; | 4.3.2.2 a | The procedure(s) adopted shall ensure that authorized editions of appropriate documents are available at all locations where operations essential to the effective The procedure(s) adopted shall ensure that functioning of the laboratory are performed; | Non-Accredited | Documented procedures must address this clause.
8.3.2 e | e) documents are uniquely identified | 4.3.2.3 | Management system documents generated by the laboratory shall be uniquely identified. Such identification shall include the date of issue and/or revision identification, page numbering, the total number of pages or a mark to signify the end of the document, and the issuing authority(ies). | Non-Accredited | Documented procedures must address this clause.
8.3.2 f | f) the unintended use of obsolete documents is prevented, and suitable identification is applied to them if they are retained for any purpose. | 4.3.2.2 c | The procedure(s) adopted shall ensure that invalid or obsolete documents are promptly removed from all points of issue or use, or otherwise assured against unintended use; | Non-Accredited | Documented procedures must address this clause.
8.3.2 f | f) the unintended use of obsolete documents is prevented, and suitable identification is applied to them if they are retained for any purpose. | 4.3.2.2 d | The procedure(s) adopted shall ensure that obsolete documents retained for either legal or knowledge preservation purposes are suitably marked. | Non-Accredited | Documented procedures must address this clause. Typical examples of "suitably marked" might be "Obsolete", "For Reference Only", "Historical Reference Only", etc.
8.4 | Control of records (Option A)
8.4.1 | The laboratory shall establish and retain legible records to demonstrate fulfilment of the requirements in this document. | 4.13.1.1 | All records shall be legible and shall be stored and retained in such a way that they are readily retrievable in facilities that provide a suitable environment to prevent damage or deterioration and to prevent loss. | Non-Accredited | Documented procedures must address the management and technical records for this standard including test datasheets, original records, cal certs, training records, purchase orders, corrective action, nonconformity results, software validation, etc.
8.4.2 | The laboratory shall implement the controls needed for the identification, storage, protection, back-up, archive, retrieval, retention time, and disposal of its records. | 4.13.1.2 | The laboratory shall establish and maintain procedures for identification, collection, indexing, access, filing, storage, maintenance and disposal of quality and technical records. Quality records shall include reports from internal audits and management reviews as well as records of corrective and preventive actions. | Non-Accredited | Management and technical records must be kept for a minimum of 5 years to ensure the ongoing validity and repeatability of the data from the date of the signature of the authorized signatory on the data package.
8.4.2 | The laboratory shall retain records for a period consistent with its contractual obligations. | 4.13.1.2 | Retention times of records shall be established. | Non-Accredited | Documented procedures must address this clause.
8.4.2 | Access to these records shall be consistent with the confidentiality commitments, and records shall be readily available. | 4.13.1.3 | All records shall be held secure and in confidence. | Non-Accredited | Documented procedures must address this clause.
8.4.2 | The laboratory shall implement the controls needed for the identification, storage, protection, back-up, archive, retrieval, retention time, and disposal of its records. | 4.13.1.4 | The laboratory shall have procedures to protect and back up records stored electronically and to prevent unauthorized access to or amendment of these records. | Non-Accredited | Documented procedures must address this clause.
8.5 | Actions to address risks and opportunities (Option A)
8.5.1 | The laboratory shall consider the risks and opportunities associated with the laboratory activities in order to: |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.1 a | a) give assurance that the management system achieves its intended results; | 4.12.1 | Needed improvements and potential sources of nonconformities, either technical or concerning the management system, shall be identified. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.1 b | b) enhance opportunities to achieve the purpose and objectives of the laboratory; |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.1 c | c) prevent, or reduce, undesired impacts and potential failures in the laboratory activities; | 4.12.1 | When improvement opportunities are identified or if preventive action is required, action plans shall be developed, implemented and monitored to reduce the likelihood of the occurrence of such nonconformities and to take advantage of the opportunities for improvement. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.1 d | d) achieve improvement. | 4.12.1 | When improvement opportunities are identified or if preventive action is required, action plans shall be developed, implemented and monitored to reduce the likelihood of the occurrence of such nonconformities and to take advantage of the opportunities for improvement. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.2 | The laboratory shall plan: |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.2 a | a) actions to address these risks and opportunities; | 4.12.2 | Procedures for preventive actions shall include the initiation of such actions and application of controls to ensure that they are effective. NOTE 1 Preventive action is a pro-active process to identify opportunities for improvement rather than a reaction to the identification of problems or complaints. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.2 b | b) how to: — integrate and implement these actions into its management system; — evaluate the effectiveness of these actions. NOTE Although this document specifies that the laboratory plans actions to address risks, there is no requirement for formal methods for risk management or a documented risk management process. Laboratories can decide whether or not to develop a more extensive risk management methodology than is required by this document, e.g. through the application of other guidance or standards. | 4.12.2 | Procedures for preventive actions shall include the initiation of such actions and application of controls to ensure that they are effective. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.5.3 | Actions taken to address risks and opportunities shall be proportional to the potential impact on the validity of laboratory results. NOTE 1 Options to address risks can include identifying and avoiding threats, taking risk in order to pursue an opportunity, eliminating the risk source, changing the likelihood or consequences, sharing the risk, or retaining risk by informed decision. NOTE 2 Opportunities can lead to expanding the scope of the laboratory activities, addressing new customers, using new technology and other possibilities to address customer needs. |  |  | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.6 | Improvement (Option A)
8.6.1 | The laboratory shall identify and select opportunities for improvement and implement any necessary actions. NOTE Opportunities for improvement can be identified through the review of the operational procedures, the use of the policies, overall objectives, audit results, corrective actions, management review, suggestions from personnel, risk assessment, analysis of data, and proficiency testing results. | 4.10 | Improvement The laboratory shall continually improve the effectiveness of its management system through the use of the quality policy, ,quality objectives, audit results analysis of data, corrective and preventive actions and management review. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.6.2 | The laboratory shall seek feedback, both positive and negative, from its customers. The feedback shall be analysed and used to improve the management system, laboratory activities and customer service. NOTE Examples of the types of feedback include customer satisfaction surveys, communication records and review of reports with customers. | 4.7.2 | The laboratory shall seek feedback, both positive and negative, from its customers. The feedback shall be used and analyzed to improve the management system, testing and calibration activities and customer service. | No | The DAP agreement along with the annual audit of the test site by qualified UL staff addresses the issues covered by these clauses. UL as the customer of the laboratory takes direct actions to satisfy itself regarding risk & improvement.
8.7 | Corrective actions (Option A)
8.7.1 | When a nonconformity occurs, the laboratory shall: | 4.11.1 | General The laboratory shall establish a policy and a procedure and shall designate appropriate authorities for implementing corrective action when nonconforming work or departures from the policies and procedures in the management system or technical operations have been identified. | Non-Accredited | Compliance is determined by review of records
8.7.1 a | a) react to the nonconformity and, as applicable:— take action to control and correct it;— address the consequences; | 4.11.3 | Where corrective action is needed, identify potential corrective actions. It shall select and implement the action(s) most likely to eliminate the problem and to prevent recurrence. | Non-Accredited | Compliance is determined by review of records
8.7.1 b | b) evaluate the need for action to eliminate the cause(s) of the nonconformity, in order that it does not recur or occur elsewhere, by:— reviewing and analysing the nonconformity;— determining the causes of the nonconformity;— determining if similar nonconformities exist, or could potentially occur; | 4.11.2 | Cause analysis The procedure for corrective action shall start with an investigation to determine the root cause(s) of the problem. NOTE Cause analysis is the key and sometimes the most difficult part in the corrective action procedure. Often the root cause is not obvious and thus a careful analysis of all potential causes of the problem is required. Potential causes could include customer requirements, the samples, sample specifications, methods and procedures, staff skills and training, consumables, or equipment and its calibration. | Non-Accredited | Compliance is determined by review of records
8.7.1 b | b) evaluate the need for action to eliminate the cause(s) of the nonconformity, in order that it does not recur or occur elsewhere, by:— reviewing and analysing the nonconformity;— determining the causes of the nonconformity;— determining if similar nonconformities exist, or could potentially occur; | 4.11.3 | Where corrective action is needed, identify potential corrective actions. It shall select and implement the action(s) most likely to eliminate the problem and to prevent recurrence. | Non-Accredited | Compliance is determined by review of records
8.7.1 c | c) implement any action needed; | 4.11.3 | Where corrective action is needed, identify potential corrective actions. It shall select and implement the action(s) most likely to eliminate the problem and to prevent recurrence. | Non-Accredited | Compliance is determined by review of records
8.7.1 d | d) review the effectiveness of any corrective action taken; | 4.11.4 | Monitoring of corrective actions | Non-Accredited | Compliance is determined by review of records
8.7.1 e | e) update risks and opportunities determined during planning, if necessary; |  |  | No | Risks & Opportunities not required of DAP participants.
8.7.1 f | f) make changes to the management system, if necessary. | 4.11.3 | The laboratory shall document and implement any required changes resulting from corrective action investigations. | Non-Accredited | Compliance is determined by review of records
8.7.2 | Corrective actions shall be appropriate to the effects of the nonconformities encountered. | 4.11.3 | Corrective actions shall be to a degree appropriate to the magnitude and the risk of the problem. | Non-Accredited | Compliance is determined by review of records
8.7.3 | The laboratory shall retain records as evidence of: | 4.11.3 | The laboratory shall document and implement any required changes resulting from corrective action investigations. | Non-Accredited | Compliance is determined by review of records
8.7.3 a | a) the nature of the nonconformities, cause(s) and any subsequent actions taken; | 4.11.3 | The laboratory shall document and implement any required changes resulting from corrective action investigations. | Non-Accredited | Compliance is determined by review of records
8.7.3 b | b) the results of any corrective action. | 4.11.3 | The laboratory shall document and implement any required changes resulting from corrective action investigations. | Non-Accredited | Compliance is determined by review of records
8.8 | Internal Audits (Option A)
8.8.1 | The laboratory shall conduct internal audits at planned intervals to provide information on whether the management system: | 4.14.1 | The laboratory shall periodically, and in accordance with a predetermined schedule and procedure, conduct internal audits of its activities to verify that its operations continue to comply with the requirements of the management system and this handbook. | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.1 | a) conforms to:— the laboratory’s own requirements for its management system, including the laboratory activities;— the requirements of this document;b) is effectively implemented and maintained. | 4.14.1 | The internal audit program shall address all elements of the management system, including the testing and/or calibration activities. It is the responsibility of the quality manager to plan and organize audits as required by the schedule and requested by management. NOTE The cycle for internal auditing should normally be completed in one year. | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 | The laboratory shall: |  |  | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 a | a) plan, establish, implement and maintain an audit programme including the frequency, methods, responsibilities, planning requirements and reporting, which shall take into consideration the importance of the laboratory activities concerned, changes affecting the laboratory, and the results of previous audits; | 4.14.1 a | The laboratory shall periodically, and in accordance with a predetermined schedule and procedure, conduct internal audits of its activities to verify that its operations continue to comply with the requirements of the management system and this handbook. The internal audit program shall address all elements of the management system, including the testing and/or calibration activities. It is the responsibility of the quality manager to plan and organize audits as required by the schedule and requested by management. NOTE The cycle for internal auditing should normally be completed in one year. | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 b | b) define the audit criteria and scope for each audit; |  |  | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 c | c) ensure that the results of the audits are reported to relevant management; |  |  | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 d | d) implement appropriate correction and corrective actions without undue delay; | 4.14.2 | When audit findings cast doubt on the effectiveness of the operations or on the correctness or validity of the laboratory's test or calibration results, the laboratory shall take timely corrective action, and shall notify customers in writing if investigations show that the laboratory results may have been affected. | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.8.2 e | e) retain records as evidence of the implementation of the audit programme and the audit results. NOTE ISO 19011 provides guidance for internal audits. | 4.14.3 | The area of activity audited, the audit findings and corrective actions that arise from them shall be recorded. | No | This clause is not applicable due to the fact that participation under the DAP program requires a UL qualified auditor to assess all required elements of ISO/IEC 17025:2017 on a regular cycle to determine compliance, and continuation in the program. UL decides to what extent to audit certain areas of the management system and includes direct observation of staff performing tests to ensure higher levels of confidence that the lab will provide compliant test results.
8.9 | Management reviews (Option A)
8.9.1 | The laboratory management shall review its management system at planned intervals, in order to ensure its continuing suitability, adequacy and effectiveness, including the stated policies and objectives related to the fulfilment of this document. | 4.15.1 a | In accordance with a predetermined schedule and procedure, the laboratory's top management shall periodically conduct a review of the laboratory's management system and testing and/or calibration activities to ensure their continuing suitability and effectiveness, and to introduce necessary changes or improvements. NOTE 1 A typical period for conducting a management review is once every 12 months. NOTE 2 Results should feed into the laboratory planning system and should include the goals, objectives and action plans for the coming year. NOTE 3 A management review includes consideration of related subjects at regular management meetings. | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 | The inputs to management review shall be recorded and shall include information related to the following: | 4.15.1 b | The review shall take account of: | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 a | a) changes in internal and external issues that are relevant to the laboratory; |  |  | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 b | b) fulfilment of objectives; |  |  | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 c | c) suitability of policies and procedures; | 4.15.1 a | the suitability of policies and procedures; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 d | d) status of actions from previous management reviews; |  |  | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 e | e) outcome of recent internal audits; | 4.15.1 c | the outcome of recent internal audits; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 f | f) corrective actions; | 4.15.1 d | corrective and preventive actions; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 g | g) assessments by external bodies; | 4.15.1 e | assessments by external bodies; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 h | h) changes in the volume and type of the work or in the range of laboratory activities; | 4.15.1 g | changes in the volume and type of the work; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 i | i) customer and personnel feedback; | 4.15.1 h | customer feedback; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 j | j) complaints | 4.15.1 i | complaints; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 k | k) effectiveness of any implemented improvements; | 4.15.1 j | recommendations for improvement; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 l | l) adequacy of resources; |  |  | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 m | m) results of risk identification; |  |  | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 n | n) outcomes of the assurance of the validity of results; | 4.15.1 f | the results of interlaboratory comparisons or proficiency tests; | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.2 o | o) other relevant factors, such as monitoring activities and training. | 4.15.1 k | other relevant factors, such as quality control activities, resources and staff training. | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
8.9.3 | The outputs from the management review shall record all decisions and actions related to at least: a) the effectiveness of the management system and its processes; b) improvement of the laboratory activities related to the fulfilment of the requirements of this document; c) provision of required resources; d) any need for change. | 4.15.2 | Findings from management reviews and the actions that arise from them shall be recorded. The management shall ensure that those actions are carried out within an appropriate and agreed timescale. | No | This clause is not applicable due UL’s qualified auditors assessing the suitability and effectiveness of the laboratory’s management system on an annual cycle.
End of New Version